FLUISPOTTER

Device Specification Repository

Version 2.29  |  2026-07-23

CONFIDENTIAL — Internal Use Only

Fluispotter ApS

sten@fluispotter.com

Contents

1. Document Overview

2. Motor Specifications — Faulhaber 1512 SR

2.1 Motor Family

2.2 Electrical Characteristics by Voltage Variant

2.3 Gearbox Options

2.4 Selected Configuration: 324 : 1 Gearbox

2.5 Encoder

2.6 Connector Pinout

2.7 Mechanical Mounting

3. Battery & Power System

3.1 Primary Cells — 2× Saft LS14500

3.2 Power Architecture

3.3 Runtime Estimate

3.4 ULP Power Budget — Worst-Case Estimate (5-Day Session)

4. Sensors & Electronics

4.1 VEML Light Sensor — Line Spot

4.2 Optical Encoder (Integrated — see Motor section)

4.3 Microcontroller / Control Electronics

4.4 Temperature and Humidity Sensor

4.5 Blood Pump Fluid Sensor (Blood / Air Detection)

4.6 AC Pump Fluid Sensor (Anticoagulant / Air Detection)

4.7 Valve Position Detection

4.8 Hall / Reed ON Switch Mechanism

4.9 Indicator LED

4A. Development Board (C03191) — Electronics Reference

4A.1 Board Overview

4A.2 Revision History

4A.3 Key Component Summary

4A.4 Hall Effect Sensors

4A.5 ESP32-S3 GPIO Pin Assignment

4A.6 Board Configuration Constants (firmware board_config.py)

5. Mechanical & Pump

5.1 Blood Pump — Stroke Volume

5.1.1 Blood Pump Summary

5.1.2 Onshape Parametric Variables

5.2 Drive Train Summary

5.3 Housing & Structural

5.4 Anticoagulant (AC) Pump

5.4.1 AC Pump Summary

5.4.2 Onshape Parametric Variables

5.5 ULP-RISC-V Pump Speed Controller (Simple-P 16-tick)

5.6 Hall-Wake Session Architecture

5.6.1 Light-Sleep State Map (Power Management — SWR-243..250)

5.6.2 Session Recovery (Power-Loss Resume)

5.7 Blood Sampling Process

5.8 Fatal Error Messages

6. Synchronize Service — Raspberry Pi Hub Server

6.1 Overview & Architecture

6.2 Running the Service

6.3 API Endpoints

6.4 Database Schema

6.5 Device Communication Protocol

6.6 WiFi Connectivity Tracking

6.7 Web UI

6.8 MicroPython Client

7. Software Architecture — Component Overview

7.1 System Diagram

7.2 Component 1 — Monitor App (macOS)

7.2.1 Passive Telemetry Mode (Development Only)

7.3 Component 2 — Control App (macOS)

7.4 Component 3 — Synchronize Service (Raspberry Pi)

7.5 Component 4 — ESP32-S3 Main Firmware (MicroPython)

7.6 Component 5 — ULP Co-processor Firmware (C + RISC-V Assembly)

7.7 Repository Structure

7.8 Flashing the Firmware Image (MicroPython)

7.9 Building a New Firmware Image — the Verified Method

8. Open Items & Design Log

8.1 Open Items

8.2 Session Log

9. Source Documents & References

1. Document Overview

This document is the living specification repository for the Fluispotter mechanical blood-pump device. It consolidates hardware datasheets, engineering calculations, component pinouts, and design decisions in a single reference. All sections should be updated as the design evolves.

Document title

Fluispotter Device Specification Repository

Version

2.29

Date

2026-07-20

Author

Sten Velschow — sten@fluispotter.com

Status

Active / In Development

Classification

Confidential — Internal Use Only

Change Log

Version

Date

Author

Changes

1.0

2026-03-30

Claude

Initial release. DevBoard C03191 (Rev A–E) indexed; MCU, sensors, GPIO pinout, Hall sensors confirmed. Blood pump and AC pump parametric data added (Sections 5.1, 5.4). Power budget and runtime estimates added.

1.1

2026-03-30

Claude

§3.1/3.2: Battery count corrected to 2× Saft LS14500; electronics supply voltage row added (3.3 VDC, LDO LD39200DPUR). §2.2: Motor voltage variant confirmed as 12 VDC (Faulhaber 1512 U 012 SR). §2.5: Encoder XOR/tacho signal description added.

1.2

2026-04-11

Claude

§4.4–4.8 added: Temperature & Humidity Sensor (SHT40-AD1B-R2), Blood Pump Fluid Sensor, AC Pump Fluid Sensor, Valve Position Detection, Hall/Reed ON Switch Mechanism.

1.3

2026-03-31

Claude

§4.8 corrected: Hall/Reed ON switch mechanism operates opposite to previous description. Packaging magnet holds device in hard power-OFF; removing device powers on.

1.4

2026-03-31

Claude

§3.3 power budget rewritten for 2× Saft LS14500 (18.72 Wh raw, 16.85 Wh usable at 90% converter efficiency). Full power budget table added. Total ~36.5 mW. Revised runtime: 462 h ≈ 19.3 days.

1.5

2026-04-11

Claude Code

§4.3: ESP32-S3 sleep modes documented with power consumption and wake latency. CPU frequency clarified. §3.4: ULP worst-case 5-day power budget. §5.1.3: measured DAC sweep chart. §5.5: ULP-RISC-V autonomous pump speed controller documented.

1.6

2026-04-11

Claude Code

§5.5 rewritten: ULP now includes hall switch edge detection during BP_RUNNING — detects 1→0 edges, increments bp_hall_revolutions, fires wakeup_main_cpu(). New shared variables: bp_hall_revolutions, main_cpu_state. §5.6 added: Hall-wake session architecture. §4.3 updated: Light Sleep confirmed as primary session mode. §6 added: Synchronize Service (Raspberry Pi hub server).

1.7

2026-04-11

Cowork

§7 added: Software Architecture — Component Overview. Documents all five software components: Monitor app (macOS, PyQt6), Control app (macOS, PyQt6), Synchronize Service (RPI, aiohttp), ESP32-S3 main firmware (MicroPython), and ULP co-processor firmware (C + RISC-V asm). Includes system diagram, repository structure, and inter-component communication. Previous §7/8 renumbered to §8/9.

1.8

2026-04-11

Cowork

§7.2.1 added: Passive Telemetry Mode. Firmware background task emits JSON telemetry on stdout, gated by /telemetry flag file. Monitor app switched from raw-REPL injection to passive serial listen — no longer interrupts the running state machine.

1.9

2026-04-11

Cowork

§5.5 rewritten: PID gain scaling clarified. Gains are raw integers written directly to ULP memory (not ×1000 encoded). Per-revolution µs error scaling documented with full derivation. Kp tuning table added (measured Kp=1–32 at SPR=16 s). Firmware boot default changed from Kp=1 to Kp=32 (stable, ~0 steady-state error). Property wrappers in pump_regulator.py fixed to match raw integer semantics.

2.0

2026-04-11

Cowork

§5.7 added: Blood Sampling Process. Full 8-step dose collection sequence with exact encoder pulses (prepare 720, collect 3 067) and timeouts. Flush cycle (2+8 min), session loop, deque confirmation logic, ULP park/drive integration, and state transition diagram documented.

2.1

2026-04-11

Cowork

§5.7.6 rewritten: Dose collection switched from sequential await chain to tick-based state machine (50 ms tick). Prepare and collect use ULP-managed dosing (bp_dosing/bp_dose_target) instead of CPU-side pulse counting. Tape smear and valve close now run concurrently; BP restarts immediately on valve close. §5.7.7 updated: documents dose() ULP integration replacing drive().

2.2

2026-04-12

Claude Code

§7.6 updated: ULP firmware now carries its own version constant (ulp_version), independent of the MicroPython firmware version. Packed as (major<<16 | minor<<8 | patch) in shared memory; readable via mem32 and included in telemetry as ulp_ver. New deploy_ulp.py script automates the full 13-step ULP build/deploy/verify cycle: build, address regeneration, atomic remap of consumer files, version stamping, binary + Python deployment via mpremote, stale .mpy cleanup, state wipe, board reset, RTC memory verification, version reporting, and telemetry sanity checks.

2.3

2026-04-12

Claude Code

§6.7: Web UI URL updated to canonical address http://fluispotter.internal/synchronize/ui/devices.

2.4

2026-04-12

Claude Code

§4.9 added: Indicator LED. GPIO 44, binary ON/OFF. Full state-to-LED mapping table with blink counts (1–6) indicating preparation workflow step. Fatal error rapid-blink pattern documented.

2.5

2026-04-12

Claude Code

§5.8 added: Fatal Error Messages. Documents all 7 fatal timeout error codes with human-readable messages logged to event store and fatal-error.txt. Covers valve position detection (3 codes) and blood pump motor/encoder/hall failures (4 codes).

2.6

2026-04-13

Claude Code

§5.7.4–5.7.8 rewritten: Unified session loop replaces two-state SESSION/SAMPLE_BLOOD design. Single _state_session with three sub-states (IDLE, TRIGGER, DOSING) and 8-step _dosing_tick. State 3 (AC_PUMP_UPKEEP) and state 11 (SAMPLE_BLOOD) removed. ENSURE_CLOSED step added before PREPARE in the dosing sequence. Dosing extracted to _dosing_init() + _dosing_tick(ctx) for code reuse across multi-dose samples. State transition diagram updated.

2.7

2026-06-04

Claude Code

§5.7.7 updated: parking/calibration now verify motor rotation by encoder ticks (~1 revolution = PARK_ENC_TICK_LIMIT 5200) instead of a 25 s wall-clock timeout, with a 30 s hard safety backstop; new shared variable bp_park_enc_total exposes the rotation count for diagnostics. §5.7.9 added: Blood Pump Motor-Start Paths — enumerates the seven code paths that start the BP motor and which use PID regulation with stall recovery.

2.8

2026-06-08

Claude Code

§5.5 rewritten for ULP 1.2.4 positional integer-tick PID. Replaces the prior per-revolution-µs incremental law (which limit-cycled at slow setpoints by winding the DAC to DAC_HALT → init-gate re-kick). New law: error in raw RC slow-clock ticks over N=2 encoder edges; output is motor power (positional) with a power floor (64 ⇒ DAC ≤ 4031, never HALT) and output-saturation anti-windup (integral clamped to the output range). Bench-validated gains Kp=300, Ki=120000, Kd=0 both pumps; AC holds SPR 20 & 7 and BP SPR 16 with the DAC bounded well below HALT and no limit cycle (BUG 1 resolved). Documents tick time-base (internal RC ~136 kHz, calibrated period Q13.19), Target_ticks derivation, integer fixed-point shifts, kick+integral-seed start, and the BP Ki-sweep finding (ripple is physical, not gain-tunable).

2.9

2026-06-10

Claude Code

§5.5 rewritten for ULP_Operational 1.3.1 — Simple-P 16-tick, now the production steady-state speed-regulation law (replaces the positional integer-tick PID, which is retained only for BP dosing/parking). New law: percent-speed-error over an N=16 edge window with a turn-on skip; incremental power nudge pwr += err% · step_per_pct inside a 1 % deadband, saturate-and-continue, no integral (self-limiting, so no windup limit cycle); per-pump gain ac_p_step_per_pct=7 / bp_p_step_per_pct=3 + shared op_p_* knobs, host-tunable with baked-in defaults. Stall/no-start interventions documented: fast (~50 ms) stall detection decoupled from the measurement window; spin-up KICK escalates to full power after op_p_start_escalate_n=3 kicks; both failure types log motor_stop_events (running-stall MOTOR_STOP_* bits 0/1 + new never-started MOTOR_START_FAIL_* bits 2/3), captured even on a successful restart. %-error skip ladder with hall-revolution counting on every tick (incl. skipped). Bench-validated 2026-06-10: AC@kp7 holds ±1 % at skip level 6; BP@kp3 stall-free (kp7 oscillated + 1 recovered stall).

2.10

2026-06-11

Claude Code

§4.1 + new §4.1b: documented the VEML6035 read protocol (must write CONFIG 0x00 = 0x0300 to power on, read ALS data at reg 0x04, lux = count × 0.0512) — fixes a “no signal from spot/line” bug where the monitor scripts never configured the sensor and read the wrong register; and the analog fluid + pressure ADCs now require ATTN_11DB for the full 0–3.3 V range (the MicroPython default clipped pressure above ~0.95 V; the air thresholds are unaffected). §7.3: PID Tuner now has three modes (Operational tunes the LIVE Simple-P regulation via op_p_*; Test = legacy PID; Simple-P = test2 sandbox), added the Sensors-tab cutoff tuning, and documented the single-channel telemetry constraint (one injected script per active tab).

2.11

2026-06-11

Claude Code

New §4A.6 Board Configuration Constants: the firmware's board_config.py values (GPIO map, DAC registers, I²C address, motor constants, PWM, motor power/speed, parking, dosing, PID gains, fluid-sensor thresholds, sync) are now listed in the spec and surfaced live in a new Board Config tab added to both the Monitor and Control apps (parsed from the source file, grouped by section, so the display can never drift from the deployed config). Note the firmware pin numbers are the authoritative deployed assignment and may differ from the schematic net table in §4A.5 (e.g. the general I²C bus). Companion app specs added: docs/monitor_spec_repo.html and docs/control_spec_repo.html.

2.12

2026-06-12

Claude Code

Gear-ratio correction (found by the verification framework's first run, resolved against the Faulhaber “Gearheads reduction ratios” document, Spur Gearheads 1512…SR): the selected gearbox is nominal 324:1, exact 323.6279268:1 — PUMP_MOTOR_GEAR_RATIO was erroneously 326 and is corrected to 324 in board_config.py + fluispotter_sv.py (§4A.6 row updated; exact-ratio note added to §2.4). The working encoder constant 5184 (= 324 × 16; exact 5178) is unchanged — bench-validated, +0.115% from exact.

2.13

2026-06-15

Claude Code

§5.5 BP proportional gain default raised bp_p_step_per_pct 3 → 6 (operational ULP 1.3.1 → 1.3.2). Bench-validated 2026-06-15 on the hall-interval SPR metric (true time between hall revolutions, not the ripple-prone encoder-batch sub-sample): at step 6 the blood pump converges 16 s/rev stall-free (motor_stop_events = 0) and tighter than 3 — baked-in run hall-interval mean 15.96 s/rev; AC unchanged (20.09 s/rev, no regression). 7 still oscillates and stays excluded. Control app writes step 6 when arming BP (v1.1.9); requirements registry SWR-105/123 + doc updated to 1.3.

2.14

2026-06-15

Claude Code

§4.1 + §4A.6: VEML6035 SPOT/LINE designations corrected — they were reversed in software. The SPOT receiver is physically at I²C 0x3a and the LINE receiver at 0x29. board_config.py now defines VEML_SPOT_ADDRESS = 0x3a and VEML_LINE_ADDRESS = 0x29 (the latter was 0x3a); fluispotter.py/fluispotter_sv.py instantiate the spot receiver at the explicit 0x3a (it previously relied on the VEML6035 driver-default 0x29); the Monitor/Control sensor scripts + panels read spot from 0x3a and line from 0x29. Requirements SWR-125/223 + doc updated to 1.4; Monitor app v1.1.2, Control app v1.1.11. Also (Control v1.1.10): new digital ON/OFF sparkline on the Sensors tab for the reed/Hall power-ON switch (GPIO5).

2.15

2026-06-15

Claude Code

§4A.6: new runtime calibration overlay — 25 tunable board_config constants are overridable at boot from /board_overrides.json WITHOUT a new firmware version (loader fluispotter/config.py; hardware constants guarded; FNV-1a config hash; a config event_store record logged at boot with hash/source/changed; /board_config.lock freeze hook). Applied once at boot (no runtime mutation); fluispotter.py reads tunables via the accessor. §4.1 fluid-threshold comments reconciled to the authoritative Blood<Air<Water/AC ordering (the active blood-detect comparison was already correct). Constants count 68→69 (the spot-address row). Requirements SWR-224…229; registry doc 1.5.

2.16

2026-06-16

Claude Code

§4.1b: documented the pressure sensor part & range — NXP MPXV5004DP differential dual-port, rated 0–3.92 kPa, ratiometric Vout = Vs·(0.2P+0.2); firmware reads raw µV (no kPa conversion). Noted the 3.3 V ADC-ceiling caveat (only ~0–2.1 kPa fits on a 5 V supply). Part + range also added as a comment in board_config.py.

2.29

2026-07-23

Claude Code

Learnings 2026-07-20→23 entered (owner instruction). §5.5: #48 RTC-pad register-clock defense (mechanism, tick-top-only DEFEND_CLOCK, stack-canary economics, TestOp dev track, op 1.16.0 gate). NEW §5.6.3: hall-deque integrity — registration proven complete (120/120 + 70/70 across 12 reboots, zero ring overflow), the three crossing-stamp corruption modes + the fw 0.1.131 fixes (wrap clamp / rate hygiene / seed) with live acceptance, sampled_at = when the blood ENTERED THE CATHETER AT THE TIP, backlog dosing pipeline + reboot 7-rev refill. §7.5: stale “no watchdog” and old SWR-235 text replaced with the current model (30 s session WDT + 120 s pre-session floor + park_wdt bench rule + serial_sink measure-and-latch and its #53 never-return hole); clock-first boot ordering (SWR-295/296, #52); NEW §7.5a boot-loop breaker + quiet-console latch + bootstat breadcrumb protocol with the crumb-id table. §7.8: CRITICAL #54 warning — IMG-4 freezes main.py, the filesystem copy never executes (proven live); entry-point rule until IMG-5. §8.y design log + session-log row for the whole arc.

2.28

2026-07-20

Claude Code

§4A.5 ESP32-S3 GPIO Pin Assignment table corrected (owner review of a planned GPIO9 hardware pull-down triggered an audit against board_config.py, the authoritative deployed source). The table had drifted from reality in two ways: (1) several rows carried a correct signal NAME on the WRONG GPIO NUMBER — the GPIO9–13 cluster had ENABLE_AC_PUMP_DC-DC sitting on GPIO11 instead of GPIO9 with the external-flash SPI signals shifted around it, and the GPIO38–44 block had MASTER_POWER/SDA_GEN/SCL_GEN/RUN_TAPE_MOTOR/RUN_INDICATION_LED all on the wrong rows; (2) GPIO1–3 still showed the spot/line-LED and valve open/close feedback assignments that §4.1 and §4.7 had already corrected to GPIO43/GPIO42/GPIO46 respectively (changelog 2.14, 2.17) — the prose was fixed at the time but this table was never brought back into sync, and GPIO42/GPIO46 were consequently missing rows entirely. All 30 assigned GPIOs re-verified 1:1 against the live PIN_* constants (confirmed no duplicate GPIO assignment exists); the rest of the document's GPIO prose (§4.1, §4.7, §7.5 DAC I²C) was already correct and required no change. Independently verified via the official ESP-IDF/Espressif ESP32-S3 pin reference that the four ESP32-S3 strapping pins are GPIO0, GPIO3, GPIO45 and GPIO46 only — GPIO9 (the AC pump enable, subject of the pull-down review) is not among them, confirming a pull-down there cannot affect boot mode, JTAG source, VDD_SPI voltage or ROM print/console selection; note added inline in the table.

2.27

2026-07-18

Claude Code

Battery endurance proof, time integrity, multi-spot findings + planned batch (fw 0.1.115, ULP 1.12.0, hub 1.13.2). §8 design log gains five dated entries: (a) 1100 mAh run-to-death RESULT — 44 h 10 m at ~24.9 mA average; death = SILENT STOP IN LIGHTSLEEP in the safe parked posture (no brownout storm, zero resets, zero data loss; early warning = a sync-timeout burst in the final 5.2 h); SWR-287 low-battery detection OPEN; (b) time-integrity chain verified + Time 0 redefined to operator submit (SWR-283..285; current receipt-based anchor measured +3 s live) and the persistent device session number (SWR-286); (c) MULTI-SPOT findings from the 3-spot bench campaign: intra-set free-run restart (~1 wasted rev + skipped tube segments, SWR-288), hall crossings counted ONLY in BP_RUNNING — dose/park crossings invisible to the blood-segment ledger (SWR-289: count in every state + ULP-owned crossing-timestamp ring), spots-vs-timepoints reporting (SWR-290); (d) pre-session sample on command in AWAIT_SESSION (SWR-291, shared routine with SWR-280); (e) session reports are dual-signed on the hub (detached OpenSSH + embedded certificate under the 20-year Fluispotter Root CA) — hub spec §11. Registries: SW 1.73 (157 reqs), HUBR 1.14 (32 reqs).

2.26

2026-07-16

Claude Code

Actuator power-on safety (SWR-276, IMG-4) + battery-hardening batch (fw 0.1.109→0.1.114, ULP 1.12.0). Field fault: AC pump ran at FULL power from power application (magnet removal) until firmware claimed the pins — DAC7574 POR is zero-scale = full drive in the inverted convention, and the AC driver enable is biased ON board-side while GPIO9/17 both FLOAT at reset (S3 datasheet v1.6 Table 2-1 — no internal pulls; the pumps differ only in driver-side bias). THREE software layers now command actuators OFF: second-stage bootloader hook (~tens of ms, IMG-4), boot.py (~0.5 s, field-measured), main.py; hardware pull-down on the AC enable queued for the next HW revision. Also: ULP 1.12.0 cooperative AC service through BP dose/coast/park (SWR-271); run-monitor keys off ULP hall counters (SWR-270); shipping-magnet reed = complete power-off + lightsleep wake source (SWR-272); no session restore before AC bag punctured (SWR-273); no lightsleep during hub sync — the field WDT class fix (SWR-274); ULP in-place self-recovery without power-cycle + stack low-water instrument (SWR-275); image builds now PIN managed components (tinyusb ==0.19.0~3) — an unpinned registry drift broke reproducibility. §4.8, §5.5, §5.6, §7.5, §7.9, §7.10 updated.

2.25

2026-07-13

Claude Code

BP Simple-P bench defaults adopted (owner PID-Tuner values) — ULP 1.10.0, fw 0.1.108, SWR-102..108 rewritten. The shared op_p_* knobs are SPLIT PER PUMP so BP tuning cannot move the AC pump: BP kp 10 (was 6), max-err clamp 50% (was 25), power floor 20 (was 64), skip ladder max level 9 / climb-after 2 good (was 6/4), and a bench-tuned BP ladder {0,4,8,16,24,64,128,256,512,1024}; the AC pump keeps ALL previous values including POWER_FLOOR=64 and 2^idx pacing (the SWR-124 slow-setpoint guard). SPR 16 s, start power 3200, N=16, max power 4095 were already the defaults. §5.5 values updated accordingly. Also: deploy_ulp.py no longer auto-stamps version.py (it overwrote fw 0.1.107 with a legacy 0.1.73 identity on-board during this change — live incident, repaired by 0.1.108).

2.24

2026-07-13

Claude Code

Firmware 0.1.102–0.1.106 feature batch + new §7.10 Versioning & Image Provenance. Three new hub commands (§7.4 command list now 13): clean_event_store (SWR-263 — wipes + re-initializes the on-device event store UID-CONTINUOUSLY via a persisted uid base in /event_uid_base, crash-safe two-phase (index removal = commit point, boot finishes/rolls back), idempotent per command uid, ack written into the FRESH store; born from the 2026-07-13 store-full incident: an unconcluded FINISH logged 'WiFi disconnected' once a minute for 37 h and filled the 8192-entry store — hub-mute + brick-on-next-boot), and telemetry_on/telemetry_off (SWR-267 — /telemetry flag + live task start/stop, default OFF). SWR-264: the boot config append degrades to a log line at store cap (boot can never brick on a full store again). SWR-265: post-dose serialization — the dead-point kick settles (RAM-only ULP telemetry poll, 3 s cap) BEFORE the sync’s flash commit + Wi-Fi TX ("do one, then the other"); both 42-run WDT(3) resets had fired in exactly that overlap; field-proven same day 20/20 zero resets. SWR-266: hub-not-found behavior — ONE stored event per outage (scan-classified 'Hub not found' vs WiFi-down), retries serial-only, events cache locally, recovery event on next sync; known-absent hub fast-fails in ~2 s instead of the 30 s blind association timeout. SWR-268 / §7.10: IMG registry (verification/images.py, IMG-0..3), boot config event carries the runtime fork-commit fingerprint, CI gate pins the runtime gitlink to a registered buildable/active image. C-level CDC fix registered as IMG-3 (mp_usbd_cdc_tx_strn zero-progress break, fork e9d9c06, image tag v0.1.107). Hub service 1.3.0–1.5.0 cross-refs: command buttons, ingestion hardening (HUBR-309/310).

2.23

2026-07-09

Claude Code

New §7.9 Building a New Firmware Image — the Verified Method: the exact 9-step procedure executed and boot-proven the same day (backup → fork branch off the proven commit pin with ONE sdkconfig commit → gitlink bump only → v* tag CI build → pre-flash gates (runtime-version string match + header/hash check — gate 1 alone would have stopped v0.1.98) → Boot1 download-mode flash → plain-power-cycle boot-proof (INITIAL LED + USB 0x4001 + hub sync) → Image-0-first sequencing → backup rollback). Executed instances recorded: v0.1.99-base (runtime v1.25.0-7.g1d140090, pipeline proof) and v0.1.99 (v1.25.0-8.g3c8d3909, + CONFIG_ESP_INT_WDT_TIMEOUT_MS=1000 only) — the issue #44 interrupt-watchdog fix, delivered and boot-proven. Root cause of the v0.1.98 no-boot documented: the fork-tip submodule bump wholesale-swapped the runtime to MicroPython v1.26.0-preview.151, which hard-hangs in early C startup before USB; both intended sdkconfig lines were exonerated (BOD LVL_SEL_7 = 2.44 V is the most-permissive esp-idf S3 default; the INT_WDT arms only after the FreeRTOS scheduler starts). Brownout deferred.

2.22

2026-07-09

Claude Code

Corrected §7.8 after the v0.1.98 rebuilt image was found NOT to boot and was rolled back. Removed the incorrect PPK2/brownout “power” claim — the no-INITIAL-LED + spot-LEDs-on + no-USB signature is a non-booting image, not inadequate supply and not a dead board (standing rule: it is never the hardware); the PPK2 note is retained only as an operational “don’t confuse the PPK2 port with the board / don’t cut DUT power” item. Corrected the ROM download-mode procedure to the verified method (physical Boot1=GPIO0 held through power-on = app-independent download mode, PID 0x1001; this board has no reset/EN button; esptool --after hard-reset RTS does not restart it → hand power-cycle to boot; running app enumerates as native USB-CDC PID 0x4001). Reframed the sdkconfig paragraph: v0.1.98 stacked INT_WDT + brownout on a year-newer runtime and would not boot — the fix is being re-delivered one sdkconfig change per image (interrupt-watchdog first, brownout last), each versioned + boot-proven on the bench, with every separately-deliverable unit carrying its own build-provenance stamp.

2.21

2026-07-09

Claude Code

New §7.8 Flashing the Firmware Image (MicroPython): the three deployment layers (MicroPython image via esptool / loose .py app via serial raw-REPL / ULP binary via deploy_ulp.py), the CI build path (a v* git tag → build.yml → fluispotter-<tag>.bin release asset, cloned via the SUBMODULE_TOKEN PAT), the write-flash 0x0 procedure (overwrites bootloader+partitions+app but preserves the littlefs FS at 0x200000 so the .py/ULP/config/state survive a runtime swap), the ROM download-mode entry the S3 native-USB requires (machine.bootloader() from the REPL / boot-catch / physical BOOT+RESET — esptool alone cannot enter it and wedges the board), and full-image backup/restore via read-flash. Also included initial flashing power/recovery guidance and an 0.1.98 sdkconfig summary — both corrected in 2.22 (the PPK2/brownout “power” claim was wrong; the no-boot was a non-booting rebuilt image, not inadequate supply).

2.20

2026-07-08

Claude Code

New §5.6.2 Session Recovery (Power-Loss Resume) (fw 0.1.92/0.1.93, SWR-257/258): persisted state (/state.dat + /time-table.dat), the 7-step boot algorithm (valve close → restore → one-shot hub reconnect/NTP → ULP health verify with AWAIT_CALIBRATION fallback → counters/schedule → AC cycle + explicit yield → SESSION/FINISH), and the AC-first invariant at every BP start. Documents the two 2026-07-08 incidents that drove it: a power cut leaves RTC memory as noise which fooled the old INITIALIZED != 0 start gate (session resumed onto a never-running ULP: no regulation, parks unserviced, latched DAC), and the synchronous entry-to-lightsleep path that starved the AC cycle task (BP ran 20-60 s unanticoagulated). Plus the stale-schedule dose-backlog operational note. §5.6.1 unchanged.

2.19

2026-07-07

Claude Code

§5.6.1 rewritten for the final flush-cycle sleep model (fw 0.1.86, SWR-243..250): the blood pump no longer runs continuously through a session — between samples it runs the 2-min-run / 8-min-park flush cycle (SWR-243; continuous pumping was a serious battery bug — reinstates the pre-4cc92da behaviour), preloads continuously at 16 SPR when a sample is <5 min out (SWR-244), and the flush cycle continues in FINISH so the line stays flushed after the last sample (SWR-248). The CPU light-sleeps in every phase except dosing and fault handling (SWR-245): the flush RUN phase and the preload/loading (TRIGGER) window sleep between hall ticks — the ULP wakes the CPU per hall edge to update the length-7 deque (SWR-247) — and the 8-min REST (motor parked + off) measures ~20 mA on the bench. Sleepers are mutually exclusive by construction (flush cycle in IDLE/FINISH vs session loop in TRIGGER). Hub sync is batched to flush RUN boundaries + pre/post-dose + faults (SWR-246, supersedes per-hall-tick SWR-242; SWR-240 superseded by SWR-245). CALIBRATION entry now verifies ULP health (stack-overflow canary clear + core loop advancing, SWR-250) and the tripped canary (#36 silent pump-dead ULP, RTC-latched, power-cycle-only) is reported as an error once per boot (SWR-249) — after a canary trip caused a silent calibration no-start misread as a mechanical fault. OI-12 largely confirmed: ULP wake from machine.lightsleep verified, ~20 mA REST floor measured; remaining = RUN/preload PPK2 profile under 0.1.86.

2.18

2026-07-07

Claude Code

New §5.6.1 Light-Sleep State Map (SWR-240 — Option A): per-state table of where machine.lightsleep() is enabled and what wakes it (fw 0.1.80, built on the proven 0.1.79 base). The design is single-owner: only the main session/FINISH loop sleeps the CPU via idle_lightsleep(backstop) — machine.lightsleep(LIGHTSLEEP_BACKSTOP_MS=60 s), woken EARLY by the ULP hall/stall/fail interrupt, bounded by the RTC backstop. The multi-sleeper timed election was removed entirely (it hung CALIBRATION and left the tape reel energized). Other tasks (telemetry, AC/BP cycles) use plain asyncio and never sleep the CPU; DOSING/BP-ACTIVE stay awake; pre-session runs no shared coordinator (the incident invariant). RTC deadlines (sample schedule, FINISH heartbeat, BP idle) survive naps; the length-7 loading deque is count-based; fault-confirm grace is inhibit-protected; a wake-side log marks each nap. New OI-12: bench-confirm (PPK2) whether the ULP interrupt cuts the nap short and the resulting idle power vs the §5.6/§3.3 model.

2.17

2026-06-19

Claude Code

Spec-vs-implementation audit corrections (264/330 claims matched; this fixes the confirmed factual + stale items). §4.7 valve-position GPIOs corrected: SIGNAL_VALVE_OPEN/CLOSE were GPIO2/GPIO3 → actual PIN_IN_VALVE_OPEN = 42 / PIN_IN_VALVE_CLOSE = 46 (GPIO2 = DAC I²C SCL, GPIO3 = pressure ADC). §4.1 spot/line emitter supply corrected to GPIO43 (GPIO48 powers the fluid receivers). §7.6 ULP firmware: dac7574.c uses the ESP-IDF ULP-RISC-V RTC-I2C hardware peripheral (NOT bit-banged; the ULP does have hardware I²C); pulse_counter.S/variables.c/config.h are legacy stubs not in the build; documented the three binaries (operational/test/test2 via ULP_TEST_BUILD/ULP_TEST2_BUILD), the 8176-byte RTC reserve and 4592-byte runtime-swap ceiling; deploy_ulp.py is 14 steps (step 4 = import-verification, not address remap; step 11 = full .text+.rodata verification). §7.5 main firmware: added the deliberate no-watchdog design and non-blocking serial (drop frames when USB-CDC not draining), the enumerated INITIAL…FINISH state machine + async background tasks, the 1 s poll / full-sync-every-5th cadence; removed the stale fixed ULP base address (addresses are regenerated per build — see variables.py). §6.4: removed the non-existent maintain_anticoagulant_pump command (10 CommandName members), added event types settings/initialization/recovery, added GET /devices/{uid}/files. §5.1.3/§5.2/§5.6 SPR set to the board_config default 16 s/rev (overridable; was 17). §5.7: tape-reel timeout documented at the code value 5000 ms (⚠ under review vs the prior 20 s — §8.1), flush-cycle park backstop 34 s (was 480 s), 5184 = 324×16 note corrected (was 326). §5.8: the three state-transition valve timeouts (valve_close_init/valve_open_ac_bag/valve_close_catheter) are fatal and now listed. §3.3: power budget re-derived for the actual 240 MHz CPU (~165 mW / 50 mA active; runtime ~17.7 days; the motor still dominates) — the firmware does not downclock; optional 80 MHz noted (OI-11 resolved). §4A.6: 12 sections (not 11); DAC7574 0x4C device address noted. (NOTE: the ULP spin-up / dose-stall start logic in §5.5/§5.7 is being revised to an adaptive kick-start — pending a separate update.)

Sync serialization (fw 0.1.111, SWR-274): the CPU NEVER enters machine.lightsleep while an event-store synchronization is in flight, and naps stay out of the 500 ms tail after a sync completes. All three field WDT(code 3) resets died at the same checkpoint — the post-dose window where the restarted flush cycle’s nap ran concurrent with the sync’s Wi-Fi TX + flash I/O (the live-reproduced INT_WDT trigger). Same “do one, then the other” doctrine as SWR-265.

Restore gate (SWR-273, owner requirement 2026-07-16): session state is restored ONLY when the persisted state is SESSION or later. Power loss in ANY earlier state — in particular before anticoagulant_bag_punctured — boots from the top (INITIAL → AWAIT_CALIBRATION); the interrupted pre-session flow is never resumed.

Reproducibility rule (2026-07-16): managed IDF components MUST be pinned to exact versions in ports/esp32/main/idf_component.yml (espressif/tinyusb: ==0.19.0~3 — the version the proven IMG-3 build resolved). An unpinned range let the component registry drift to tinyusb 0.21.0, which renamed CFG_TUD_CDC_EP_BUFSIZE and broke the build three days after an identical build succeeded. Version pinning is part of the one-change-per-image guarantee.

2. Motor Specifications — Faulhaber 1512 SR

2.1 Motor Family

Model family: FAULHABER 1512 U … SR … IE2-8
Type: DC gearmotor with integrated 2-channel optical incremental encoder (IE2-8).
Nominal voltage variants available: 3 V, 6 V, and 12 V.

2.2 Electrical Characteristics by Voltage Variant

The Fluispotter uses the 12 VDC variant of the Faulhaber 1512 SR motor (part number 1512 U 012 SR).

Parameter

3 V variant

6 V variant

12 V variant

Nominal voltage

3 V

6 V

12 V

Terminal resistance

10.4 Ω

50.5 Ω

130 Ω

No-load speed

13,400 rpm

14,300 rpm

15,500 rpm

2.3 Gearbox Options

Gear ratio

Approx. output speed (at 15,000 rpm)

Typical use

6 : 1

~2,500 rpm

High-speed output

13 : 1

~1,150 rpm

Medium speed

39 : 1

~385 rpm

Medium-low speed

112 : 1

~134 rpm

Low speed

324 : 1

~46 rpm (free-run); ~15 rpm at load

Very low speed — selected for Fluispotter

2.4 Selected Configuration: 324 : 1 Gearbox

Note: This is the gearbox configuration selected for the Fluispotter blood-pump drive.

The datasheet ratio 324:1 is the rounded nominal value. Per the Faulhaber “Gearheads reduction ratios” document (Spur Gearheads 1512…SR), the exact reduction is 323.6279268 : 1 (210 208 960 / 649 539). With the IE2-8 encoder (8 impulses × 2 channels = 16 edges per motor revolution) the exact encoder count is 16 × 323.628 ≈ 5178 edges per output revolution; the firmware's working constant 5184 (= 324 × 16) is +0.115 % from exact — negligible and bench-validated.

Gear ratio

324 : 1

Max output speed (no load)

~15 rpm

Continuous output torque

30.0 mNm

Intermittent output torque

50 mNm

Gearbox efficiency

53 %

Assembly mass

7.7 g

Actual operating speed

17 s per rotation (≈ 3.5 rpm) — via DC-DC converter voltage reduction

2.5 Encoder

The two quadrature encoder channels (A and B) are XORed together by a discrete XOR gate on the DevBoard, producing a single tacho signal. This doubles the pulse frequency compared to either channel alone (16 XOR rising edges per motor rev). Each pump has its own dedicated XOR-processed tacho line routed to the ESP32-S3. The tacho signals are handled by the ULP-RISC-V co-processor, which counts pulses in Light Sleep mode independently for each pump without waking the main cores.

Encoder type

Integrated optical incremental — IE2-8

Resolution

8 lines per motor revolution

Channels

A and B (quadrature)

Phase shift

90° between channels A and B

Duty cycle (usage)

Encoder powered 5 % of operating time; OFF for remaining 95 %

2.6 Connector Pinout

6-pin JST-style connector on the motor/encoder assembly.

Pin

Signal

Notes

1

Motor −

Motor power negative

2

Motor +

Motor power positive

3

GND

Encoder / power ground

4

VDD

Encoder supply voltage

5

Channel B

Encoder quadrature B

6

Channel A

Encoder quadrature A

2.7 Mechanical Mounting

Mounting thread

M1.6

Screw count

2 × M1.6 screws

Thread depth

2.5 mm minimum engagement

3. Battery & Power System

3.1 Primary Cells — 2× Saft LS14500

The Fluispotter device is powered by two Saft LS14500 lithium thionyl chloride (Li-SOCl₂) primary cells. Configuration (series / parallel) to be confirmed; values below are per cell.

Cell model

Saft LS14500

Chemistry

Lithium thionyl chloride (Li-SOCl₂)

Form factor

AA (14.5 mm × 50.5 mm)

Nominal voltage

3.6 V

Nominal capacity

2,600 mAh

Typical energy

9.36 Wh

Operating temperature

−60 °C to +85 °C

Self-discharge

< 1 % per year at +20 °C

Datasheet on file

Saft LS14500 cell.pdf

3.2 Power Architecture

A DC-DC step-down (buck) converter reduces the battery voltage to approximately 2 V to achieve the target motor rotation speed. The LDO (LD39200DPUR) provides a regulated 3.3 VDC rail for all electronics.

Battery voltage

3.6 V (nominal)

Electronics supply voltage

3.3 VDC (regulated by LDO LD39200DPUR from battery)

Motor drive voltage

~2 V (via DC-DC converter)

Converter type

DC-DC buck converter

Encoder supply

VDD (from converter or battery — TBD)

Encoder duty cycle

5 % on / 95 % off

Target motor power draw

~30 mW (continuous when running)

3.3 Runtime Estimate

Estimates based on 2× Saft LS14500 (18.72 Wh total raw, 16.85 Wh usable at 90 % converter efficiency). The device spends the majority of its time in Light Sleep with the ULP co-processor active; the main CPU wakes for short active periods on each hall tick.

Battery energy

Raw energy (2 cells)

2 × 3.6 V × 2.6 Ah = 18.72 Wh

Usable energy (× 0.90 converter eff.)

16.85 Wh

Battery configuration

Series (7.2 V) or parallel (3.6 V) — TBD

Power budget (average)

Consumer

Avg. power

Basis

Motor — Faulhaber 1512, DC-DC ~2 V

30 mW

Continuous (worst-case)

ESP32-S3 deep sleep + ULP (95 %)

0.35 mW

107 µA × 3.3 V

ESP32-S3 active @ 240 MHz (5 %)

8.25 mW avg.

~165 mW (50 mA × 3.3 V, §4.3 upper bound) × 0.05

ESP32-S3 average (95/5 duty)

8.6 mW

0.95 × 0.35 + 0.05 × 165

LDO + peripheral quiescent

~1 mW

LD39200 LDO, SHT40, drivers

Total average (motor continuous)

~39.6 mW

30 + 8.6 + 1

Runtime estimate

Theoretical continuous runtime

16.85 Wh ÷ 39.6 mW ≈ 425 h ≈ 17.7 days

ESP32 only (no motor)

16.85 Wh ÷ 0.35 mW ≈ 48,140 h ≈ 5.5 years

Note: These are theoretical estimates; full runtime characterisation requires bench measurement under realistic dosing conditions (OI-07). The active-CPU term reflects the actual 240 MHz operation — the firmware does not call machine.freq() — using the §4.3 upper-bound ~165 mW (50 mA) at a 5 % active duty. The motor (30 mW continuous, worst-case) still dominates, so moving from the earlier 80 MHz assumption shifts the headline runtime only ~19.3 → ~17.7 days. (A boot-time machine.freq(80_000_000) would cut the active-CPU term ~38 % if a longer field life is later required.)

3.4 ULP Power Budget — Worst-Case Estimate (5-Day Session)

Worst-case energy over 5 days (120 h) with ULP pump controller active. All figures in mAh against 2 600 mAh cell budget.

WiFi synchronisation (dominant cost): 300 mA × 1 s × 27 000 Hall ticks / 3 600 = 2 250 mAh (86 % of one cell).

AC pump (ULP idle, 100 % duty): 1.8 mAh. Blood pump (ULP regulation, 30 % duty): 0.9 mAh.

CPU + ULP housekeeping (5 % active at 80 MHz): 186 mAh.

Total: ~2 440 mAh. Fits within one 2 600 mAh cell with ~160 mAh margin. Second cell is full reserve.

Design recommendation: Reduce WiFi sync to once per N revolutions (N=100 → 22.5 mAh, total ~210 mAh, ~62 days on two cells).

v1.6 note — hall-wake architecture. The v1.6 firmware uses Light Sleep (not Deep Sleep) between hall ticks. WiFi sync is decoupled from hall ticks and configurable independently. Light Sleep current ~2 mA is higher than Deep Sleep ~7 µA but eliminates full-reboot penalty. At 17 s/rev the CPU spends >99.97 % in Light Sleep.

4. Sensors & Electronics

4.1 VEML Light Sensor — Line Spot

Two VEML6035 ambient light sensors (Vishay). Designators: SPOT and LINE. The VEML6035 native I²C address is 0x29 (fixed); the LTC4316 translator (IC15) maps the pair onto the shared bus as SPOT = 0x3a, LINE = 0x29 (corrected 2026-06-15 — the two designations were previously reversed in software). I²C on SDA_GEN (GPIO38) / SCL_GEN (GPIO43). Enable: the SPOT/LINE emitter pair is powered by PIN_OUT_ENABLE_TAPE_EMITTERS (GPIO43, shared with the tape emitter), not GPIO48 — GPIO48 (PIN_OUT_ENABLE_FLUID_RECEIVERS) enables the blood/AC fluid receivers (see §4.5/§4.6). The prior “supply via GPIO48” was a misattribution; see the §4A.6 GPIO map for the authoritative assignment.

Read protocol (must configure first). The VEML6035 boots shut down (CONFIG register 0x00, bit 0 = ALS_SD) and returns no data until configured. Any reader (driver or host script) must first write CONFIG 0x00 to power it on — the firmware uses 25 ms integration time, gain/dg normal, sensitivity HIGH = 0x0300 — then read the ALS data at register 0x04 (16-bit little-endian) and scale lux = count × 0.0512 (the lux_resolution for that config, designingveml6035.pdf). Skipping the power-on, or reading the wrong register (e.g. 0x09), returns nothing — this was the root cause of a “no signal from the spot/line sensors” bug in the Control/Monitor sensor-monitor scripts (now fixed: they configure 0x0300 and read 0x04).

4.1b Analog Fluid + Pressure Sensors (ADC)

The Blood/air receiver (ADC GPIO7), AC/air receiver (ADC GPIO6) and the pressure sensor (ADC GPIO3) are read via machine.ADC(...).read_uv(). These ADCs must be created with atten=ADC.ATTN_11DB for the full 0–3.3 V range; the MicroPython default attenuation caps at ~0.95 V and would clip any reading above that (it left the firmware pressure reading clipped while the Control App, which already set ATTN_11DB, read the full range). Because read_uv() returns the calibrated true voltage for in-range signals, the air-sensor thresholds in board_config.py (blood: air > 800 mV, AC/water 500–800, blood < 500; AC: water < 500) are unaffected by the attenuation change.

Pressure sensor part & range. The pressure transducer is an NXP MPXV5004DP — a differential, dual-port sensor rated 0–3.92 kPa (≈ 0–39.2 mbar / 0–400 mmH&sub2;O / 0–29.4 mmHg / 0–0.57 psi). Its output is ratiometric: Vout = Vs·(0.2·P[kPa] + 0.2) — sensitivity 0.2·Vs per kPa, with a zero-pressure offset of 0.2·Vs (so ≈1.0 V at rest, ≈4.9 V full-scale on a 5 V supply). The firmware does no kPa conversion — PressureSensor.read_pressure() returns the raw read_uv() microvolts (the Sensors tab shows mV). Caveat: the ESP32 ADC tops out at ~3.3 V, so on a 5 V supply (consistent with the ~1 V resting reading that was being clipped above) only roughly the lower ~0–2.1 kPa of the rated span fits under the ADC ceiling; capturing the full 3.92 kPa needs a ~3.3 V sensor supply (full-scale ≈ 3.25 V) or an output divider — confirm the pressure rail.

4.2 Optical Encoder (Integrated — see Motor section)

Faulhaber IE2-8 encoder provides rotation feedback. See Sections 2.5 and 2.6.

4.3 Microcontroller / Control Electronics

MCU: ESP32-S3-WROOM-1-N8R8 (Espressif). Dual-core Xtensa LX7, 240 MHz. 8 MB Flash, 8 MB PSRAM. Wi-Fi + BLE 5.0. U.FL antenna. Designator: IC9.

Sleep modes. Active: 20–50 mA at 80–240 MHz. Modem Sleep: RF off only — not useful for Fluispotter. Light Sleep: ~2 mA, <1 ms wake latency, RAM retained, MicroPython resumes exactly where it stopped. This is the primary operating mode during sessions — machine.lightsleep() with no timeout, waking only on ULP RTC_CNTL_SW_CPU_INT interrupt per hall-switch revolution. See §5.6. Deep Sleep: ~7–150 µA, full reboot on wake. Reserved for long idle periods.

CPU frequency: the firmware does not call machine.freq() (verified 2026-06-19), so the CPU runs at the MicroPython default 240 MHz. The §3.3 power budget is computed at this 240 MHz figure (~165 mW / 50 mA active, §4.3 upper bound). Adding a boot-time machine.freq(80_000_000) would reduce the active-CPU term ~38 % if a longer field life is later required.

Motor drivers: 2× TI DRV8835DSSR (IC7, IC8). Dev Board: C03191 Rev E (22 Jan 2026).

4.4 Temperature and Humidity Sensor

Sensirion SHT40-AD1B-R2 (IC5). I²C on SDA_GEN/SCL_GEN. Range: 0–100 % RH / −40 to +125 °C. Supply: 3.3 V rail (always on).

4.5 Blood Pump Fluid Sensor (Blood / Air Detection)

IR emitter–phototransistor pair. Blood absorbs more IR than air. Signal conditioned by TLV313IDBVR op-amp, read by ESP32 ADC.

IR emitter

Kingbright APDA3020F3C (LED2/LED3)

Phototransistor

Kingbright APDA3020P3C (Q3/Q7)

ADC signal

ANALOG_SIGNAL_BLOOD_FLUID_SENSOR — GPIO7 (ULP-accessible)

LED enable

RUN_AC_AND_BLOOD_FLUID_SENSOR_LEDS — GPIO47

Sensor power enable

ENABLE_AC_AND_BLOOD_FLUID_SENSORS — GPIO48

4.6 AC Pump Fluid Sensor (Anticoagulant / Air Detection)

Identical topology to blood sensor. ADC: ANALOG_SIGNAL_AC_FLUID_SENSOR — GPIO6 (ULP-accessible). Shared LED enable (GPIO47) and sensor power (GPIO48).

4.7 Valve Position Detection

PIN_IN_VALVE_OPEN = GPIO42 and PIN_IN_VALVE_CLOSE = GPIO46 (corrected 2026-06-19 — the prior GPIO2/GPIO3 were wrong: GPIO2 is the DAC I²C SCL, GPIO3 is the pressure-sensor ADC, both in use elsewhere). These limit switches are read by the main CPU (GPIO42/46 are not RTC-capable, so — unlike the spec's old “ULP-accessible” note — they are not visible to the ULP). Connector: Molex 530470310. Drive: GPIO8 (PWM) / GPIO45 (direction).

4.8 Hall / Reed ON Switch Mechanism

Packaging magnet holds reed switch closed = hard power-OFF. Remove from packaging → reed opens → GPIO5 triggers power-on. Rev A–D: Standex-Meder MK24-B-3-OE (SPST NC). Rev E added: TI TMAG5231B1DQDBZR Hall (SOT-23, 1.65 V). Bypass jumper for development.

Firmware behavior (SWR-272, fw 0.1.111+): at boot the firmware asserts the master-power latch (PIN_OUT_MASTER_POWER=1) to hold the main circuit on. A background task armed from firmware start watches the reed in ALL states: magnet re-engaged (device returned to packaging) → “Reed switch engaged. Shutting down device.” → the latch is released (power.off()) = COMPLETE power-off. The reed is also armed as a lightsleep WAKE source around every nap (falling edge + wake=machine.SLEEP), so power-off is immediate even mid-nap, and the fatal-error handler independently honours the reed.

4.9 Indicator LED

Single binary indicator LED on GPIO 44 (PIN_OUT_INDICATION_LED). ON/OFF only — no colour or PWM.

State

ID

LED

Notes

Initial

1

ON (solid)

Valve close, hardware init

Await Calibration (connecting)

2

ON (solid)

WiFi scanning & waiting for hub acquisition

Await Calibration (acquired)

2

Blink ×1

Acquired; waiting for prepare_session or purge_catheter

AC Pump Upkeep

3

—

Maintenance sub-state

Calibration

4

ON (solid)

Valve exercise, tape reel, BP parking

Await AC Bag

5

Blink ×2

Waiting for anticoagulant_bag_punctured

AC Bag

6

ON → Blink ×3

AC pump runs fast, then blinks waiting for anticoagulant_primed

Await Catheter

7

Blink ×4

Waiting for catheter_inserted

Catheter

8

ON + Blink ×5

BP running; blinks while waiting for blood detection / blood_primed

Await Session

9

Blink ×6

Waiting for start_session

Session (dormant)

10

OFF

Between samples; device sleeping

Sample Blood

11

OFF

Collecting blood sample

Finish

12

OFF

Session complete

Fatal error

Toggle 500 ms

Rapid blink; resets after 2 reed presses

Blink pattern: each blink = 100 ms ON + 100 ms OFF, repeated N times, then 1.6 s pause (16 × 100 ms). The blink count indicates where in the preparation workflow the device is — higher count = closer to session start. Await Calibration uses blink ×1 while waiting for prepare_session (or purge_catheter).

4A. Development Board (C03191) — Electronics Reference

4A.1 Board Overview

Primary electronics integration platform. Board files: Fluispotter/schematics/DevBoard_C03191/

4A.2 Revision History

Revision

Date

Schematic

PCB

Eagle

Gerbers

BOM

Notes

Rev A

03 Sep 2025

YES

YES

YES

YES

YES

Initial release.

Rev B

10 Sep 2025

YES

YES

YES

YES

—

Minor update.

Rev D

12 Sep 2025

YES

YES

YES

YES

YES

Incremental change.

Rev E

22 Jan 2026

—

YES

YES

YES

YES

Current. Added TMAG5231 Hall, WROOM-1-N8R8.

4A.3 Key Component Summary

Ref

Part Number

Manufacturer

Description

IC9

ESP32-S3-WROOM-1-N8R8

Espressif

Main MCU — dual-core LX7, BLE, Wi-Fi

IC7, IC8

DRV8835DSSR

Texas Instruments

Dual H-bridge motor drivers

AC/BP HALL

DRV5032AJDMRR

Texas Instruments

Digital Hall switch, open-drain

IC6

MCP79410T-I/SN

Microchip

Real-time clock (RTC), I²C

U2

DAC7574IDGSR

TI

12-bit quad DAC, I²C

IC5

SHT40-AD1B-R2

Sensirion

Humidity + temperature, I²C

Y2

ECS-.327-7-1210-TR

ECS Inc.

32.768 kHz crystal

See full BOM files for complete component list including passive components.

4A.4 Hall Effect Sensors

DRV5032AJDMRR — AC_PUMP_HALL (GPIO15) and BLOOD_PUMP_HALL (GPIO21). Open-drain, active LOW. ULP-accessible.

TMAG5231B1DQDBZR (Rev E) — Low-power Hall sensor. GPIO and role TBD.

4A.5 ESP32-S3 GPIO Pin Assignment

GPIO0–GPIO21: RTC domain — ULP-accessible. GPIO35–GPIO48: Digital only — NOT ULP-accessible.

Re-verified 2026-07-20 against board_config.py (the authoritative deployed source, §4A.6) after an ESP32-S3 strapping-pin review found this table had drifted: several rows carried the correct signal NAME on the wrong GPIO NUMBER (a scramble in the GPIO9–13 SPI-flash/AC-enable cluster and the GPIO38–44 block), GPIO1/2/3 still showed the spot/line-LED and valve-feedback assignments that §4.1/§4.7 had already corrected to GPIO43/42/46 (changelog 2.14, 2.17) without the table being updated to match, and GPIO42 and GPIO46 (the valve limit switches) were missing rows entirely. Every row below is now cross-checked 1:1 against the live PIN_* constants; no duplicate GPIO assignment exists in board_config.py.

GPIO

Signal / Net Name

ULP

Function

GPIO0

BOOT_MODE

YES

Boot strapping

GPIO1

SDA_DAC

YES

ULP I²C SDA — DAC bus (PIN_I2C_SDA_DAC)

GPIO2

SCL_DAC

YES

ULP I²C SCL — DAC bus (PIN_I2C_SCL_DAC)

GPIO3

ANALOG_SIGNAL_PRESSURE

YES

ADC — pressure sensor (PIN_ADC_PRESSURE_SENSOR)

GPIO4

SIGNAL_BLOOD_PUMP_ENCODER_XOR

YES

Blood pump encoder tacho

GPIO5

WAKE_INTERRUPT_SIGNAL_REED_STATUS

YES

Power-on trigger (reed/Hall)

GPIO6

ANALOG_SIGNAL_AC_FLUID_SENSOR

YES

ADC — AC pump fluid sensor

GPIO7

ANALOG_SIGNAL_BLOOD_FLUID_SENSOR

YES

ADC — blood pump fluid sensor

GPIO8

RUN_VALVE_MOTOR

YES

Valve motor PWM / enable

GPIO9

ENABLE_AC_PUMP_DC-DC

YES

AC pump DC-DC enable (floats at reset — SWR-276, no internal pull; not an ESP32-S3 strapping pin)

GPIO10

FLASH_CS

YES

SPI external Flash — chip select

GPIO11

FLASH_MOSI

YES

SPI external Flash — MOSI

GPIO12

FLASH_SCLK

YES

SPI external Flash — clock

GPIO13

FLASH_MISO

YES

SPI external Flash — MISO

GPIO14

ENABLE_AC_PUMP_ENCODER

YES

AC pump encoder power

GPIO15

WAKUP_INTERRUPT_SIGNAL_AC_PUMP_HALL

YES

AC pump Hall (active LOW)

GPIO16

SIGNAL_AC_PUMP_ENCODER_XOR

YES

AC pump encoder tacho

GPIO17

ENABLE_BLOOD_PUMP_DC-DC

YES

Blood pump DC-DC enable (floats at reset — SWR-276, no internal pull; not an ESP32-S3 strapping pin)

GPIO18

ENABLE_BLOOD_PUMP_ENCODER

YES

Blood pump encoder power

GPIO19–20

USB_DM/DP_ESP

YES

USB 2.0 D−/D+ (RTC when USB inactive)

GPIO21

WAKUP_INTERRUPT_SIGNAL_BLOOD_PUMP_HALL

YES

Blood pump Hall (active LOW)

GPIO38

MASTER_POWER

NO

Master power hold / control (PIN_OUT_MASTER_POWER)

GPIO39

SDA_GEN

NO

I²C SDA — general bus

GPIO40

SCL_GEN

NO

I²C SCL — general bus

GPIO41

RUN_TAPE_MOTOR

NO

Tape reel motor

GPIO42

SIGNAL_VALVE_OPEN

NO

Valve open feedback (main-CPU read; not RTC-capable)

GPIO43

RUN_SPOT_&LINE_LED / TAPE_EMITTER

NO

Spot/line optical LED enable, shared with the tape emitter (PIN_OUT_ENABLE_TAPE_EMITTERS)

GPIO44

RUN_INDICATION_LED

NO

Blue status LED

GPIO45

SET_VALVE_MOTOR_DIRECTION

NO

Valve motor direction

GPIO46

SIGNAL_VALVE_CLOSE

NO

Valve closed feedback (main-CPU read; not RTC-capable)

GPIO47

RUN_AC_AND_BLOOD_FLUID_SENSOR_LEDS

NO

Fluid sensor LED enable

GPIO48

ENABLE_AC_AND_BLOOD_FLUID_SENSORS

NO

VEML6035 sensor power

GPIO22–34 used internally by WROOM for SPI flash/PSRAM. GPIO35–37 not connected. ESP32-S3 strapping pins are GPIO0, GPIO3, GPIO45 and GPIO46 only (Espressif ESP32-S3 datasheet); none of the actuator-enable or SPI-flash pins above are strapping pins.

4A.6 Board Configuration Constants (firmware board_config.py)

The firmware's board-type configuration lives in fluispotter/board_config.py and is deployed alongside the firmware. It is the authoritative source for the pin, register, and tuning constants the firmware and ULP coprocessor actually use. The Monitor and Control apps each surface these values live in a Board Config tab — parsed from the source file and grouped by the sections below — so the displayed configuration can never drift from what is deployed. The logical pin numbers below are the deployed firmware assignment and may differ from the schematic net names in §4A.5 (e.g. the general I²C bus).

Runtime calibration overlay (firmware-independent). A subset of 25 tunable calibration constants (fluid-sensor thresholds, pump speeds + start powers, ULP PID gains, dose degrees/pulses, parking offset, sync interval) may be overridden at boot from a JSON file /board_overrides.json on the device — adjusting them does not require a new firmware version. The loader (fluispotter/config.py) merges the overlay over the compiled defaults once at boot (no runtime mutation), guards hardware constants (pins, DAC registers, I²C addresses, motor physics) against override, validates type and range, and fingerprints the effective config with an FNV-1a hash. fluispotter.py reads the tunable values through this accessor (self.cfg.NAME) and, at boot, appends one event-store record of type config carrying the hash, source (defaults/overlay/locked), override count and rejected keys, plus a changed flag computed against the last persisted hash (/board_config.hash). A /board_config.lock file freezes the overlay out (compiled defaults only) — the hook by which calibration can later be locked back into firmware. See SWR-224…229.

GPIO pin assignments

Constant

GPIO

Constant

GPIO

PIN_IN_BLOOD_PUMP_ENCODER

4

PIN_IN_BLOOD_PUMP_HALL_SWITCH

21

PIN_IN_REED_SWITCH

5

PIN_OUT_MASTER_POWER

38

PIN_ADC_AC_RECEIVER

6

PIN_I2C_SDA

39

PIN_ADC_BLOOD_RECEIVER

7

PIN_I2C_SCL

40

PIN_OUT_ENABLE_VALVE

8

PIN_OUT_ENABLE_TAPE_REEL

41

PIN_OUT_ENABLE_AC_PUMP

9

PIN_IN_VALVE_OPEN

42

PIN_SPI_CS_FLASH

10

PIN_OUT_ENABLE_TAPE_EMITTERS

43

PIN_SPI_MOSI

11

PIN_OUT_INDICATION_LED

44

PIN_SPI_SCLK

12

PIN_OUT_VALVE_DIRECTION

45

PIN_SPI_MISO

13

PIN_IN_VALVE_CLOSE

46

PIN_OUT_ENABLE_AC_PUMP_ENCODER

14

PIN_OUT_ENABLE_FLUID_EMITTERS

47

PIN_IN_AC_PUMP_HALL_SWITCH

15

PIN_OUT_ENABLE_FLUID_RECEIVERS

48

PIN_IN_AC_PUMP_ENCODER

16

PIN_ADC_PRESSURE_SENSOR

3

PIN_OUT_ENABLE_BLOOD_PUMP

17

PIN_I2C_SDA_DAC

1

PIN_OUT_ENABLE_BLOOD_PUMP_ENCODER

18

PIN_I2C_SCL_DAC

2

DAC registers, I²C address, motor & PWM constants

Constant

Value

Notes

DAC_REGISTER_AC_PUMP_SIGNAL

0

DAC7574 channel — AC pump power

DAC_REGISTER_BLOOD_PUMP_SIGNAL

1

DAC7574 channel — blood pump power

DAC_REGISTER_BLOOD_PUMP_DIRECTION

3

DAC7574 channel — blood pump direction

VEML_SPOT_ADDRESS

0x3a

VEML6035 spot sensor I²C address

VEML_LINE_ADDRESS

0x29

VEML6035 line sensor I²C address

PUMP_MOTOR_PULSE_PER_REVOLUTION

8

One encoder channel (16 total)

PUMP_MOTOR_GEAR_RATIO

324

Gearbox nominal ratio (exact 323.6279268 per Faulhaber; was erroneously 326 until 2026-06-12)

PUMP_MOTOR_POWER_HALT

4095

DAC value = motor stopped (power = 4095 − dac)

VALVE_MOTOR_DIRECTION_CLOSE / OPEN

0 / 1

Valve direction codes

PWM_MAX

65534

LEDC PWM full scale

PWM_FREQUENCY

2000

Valve/tape PWM frequency (Hz)

ULP_PULSES_PER_OUTPUT_REV

5184

326 × 16 — encoder ticks per output revolution

Motor power & speed, parking & dosing

Constant

Value

Constant

Value

BLOOD_PUMP_SECONDS_PER_REVOLUTION

16

AC_PUMP_MEDIUM_SECONDS_PER_REVOLUTION

20

BLOOD_PUMP_CALIBRATION_SECONDS_PER_REV

8

AC_PUMP_MEDIUM_START_POWER

700

BLOOD_PUMP_START_POWER

3000

AC_PUMP_SLOW_SECONDS_PER_REVOLUTION

20

AC_PUMP_FAST_SECONDS_PER_REVOLUTION

7

AC_PUMP_SLOW_START_POWER

700

AC_PUMP_FAST_START_POWER

4050

VALVE_MOTOR_POWER

100

BLOOD_PUMP_PARK_OFFSET_DEGREES

0

BLOOD_PUMP_SAMPLE_PREPARE_DEGREES

50

BLOOD_PUMP_SAMPLE_PREPARE_PULSES

720

BLOOD_PUMP_SAMPLE_COLLECT_DEGREES

213

BLOOD_PUMP_SAMPLE_COLLECT_PULSES

3067

PID gains (ULP positional integer-tick PID), fluid-sensor thresholds & sync

Constant

Value

Notes

AC_ULP_PID_KP / KI / KD

300 / 120000 / 0

AC pump regulator gains (Control App can override at runtime)

BP_ULP_PID_KP / KI / KD

300 / 120000 / 0

Blood pump regulator gains

BLOOD_SENSOR_AIR_THRESHOLD_UV

800000

Above → air (blood-air sensor)

BLOOD_SENSOR_AC_THRESHOLD_UV

500000

Above → AC fluid, below → blood

BLOOD_SENSOR_BLOOD_THRESHOLD_UV

500000

Below → blood detected

AC_SENSOR_FLUID_THRESHOLD_UV

500000

Below → AC fluid detected (AC-air sensor)

SYNCHRONIZE_INTERVAL_MS

1000

Hub polling interval (ms)

69 constants in 12 sections. (The DAC7574 device I²C address 0x4C is given in §4A.3 / §5.5, not in these tables — the DAC rows here carry the channel registers 0/1/3 and the SDA/SCL pins.) Values are shown verbatim from the source (the Board Config tab in each app renders the same listing live). The PID-gain section reflects the ULP 1.2.4 positional law; the operational regulator has since moved to the Simple-P 16-tick law (§5.5), but the firmware retains these board_config.py defaults and the Control App tunes the live regulator at runtime.

5. Mechanical & Pump

5.1 Blood Pump — Stroke Volume

Peristaltic (rotary tubing) pump. Full stroke: 18.31 µL/rev. Active stroke (315°): 16.02 µL. Tubing: 0.76 mm ID / 1.65 mm OD. Pump OD: 14.5 mm. Dosing target: 1 µL/s → 1061.5 RPM motor.

Parameter

Value

Pump type

Peristaltic (rotary)

Full stroke volume (360°)

18.31 µL

Active stroke (315°)

16.02 µL

Park angle

45°

Degrees per µL

19.66°/µL

Motor RPM at 1 µL/s

1 061.5 RPM

Encoder steps per pump revolution

5 184 rising edges

Stroke volume formula: V = ¼π² × D_centre × d_ID². See Onshape parametric variables in §5.1.2 of v1.5 for full variable table.

5.1.3 Blood Pump DAC Sweep — Power vs SPR (Measured)

Stall below DAC ~2500. Reliable from DAC ~2600 (14 SPR) to 4095 (8 SPR). PID start power 3000 is within reliable range. Firmware target 16 SPR (BLOOD_PUMP_SECONDS_PER_REVOLUTION = 16 — board_config default, overridable; was 17 in earlier revisions).

5.2 Drive Train Summary

Motor

Faulhaber 1512 U 012 SR (12 V)

Gearbox

324 : 1

Drive voltage

~2 V (DC-DC)

Output speed

~3.75 rpm (16 s/rev — board_config default, overridable; was 17)

Continuous torque

30.0 mNm

5.3 Housing & Structural

Dimensions, material, IP rating, sterilisation, regulatory pathway: all TBD.

5.4 Anticoagulant (AC) Pump

Peristaltic (rotary), 2-roller, 0.31 mm ID tubing, 9.50 mm centre tube, 324:1 gearbox. Roller loss 11.17 %. Actual stroke: 2.001 µL. Target: 0.057 µL/s. Max runtime 25 mL reservoir: 4.16 days.

See Onshape parametric variables in §5.4.2 of v1.5 for full AC pump variable table.

5.5 ULP-RISC-V Pump Speed Controller (Simple-P 16-tick)

A fully autonomous pump speed regulation loop runs on the ESP32-S3 ULP-RISC-V co-processor, implemented in C (main.c) and loaded into RTC slow memory at boot. The loop operates independently of the main Xtensa cores while the device is in Light Sleep; both the blood pump and the AC pump are regulated. As of ULP_Operational 1.3.1 the steady-state speed regulation uses a Simple-P incremental control law over a 16-edge measurement window (“Simple-P 16-tick”), replacing the positional integer-tick PID of 1.2.4. The two blood-pump special modes — precision dosing (BP_DOSING) and hall parking/calibration (BP_PARKING) — retain the legacy positional PID (run_pid()), unchanged.

Execution model. The controller is tick-based: each ULP wake runs ac_tick() then bp_tick(), performing at most one measurement plus one power step per regulated tick, with all loop state held in RTC slow RAM between wakes — not a blocking loop.

Encoder sampling (N = 16, edge-locked). The ULP enables encoder power, polls the XOR quadrature signal (GPIO4 blood / GPIO16 AC), discards ENCODER_SKIP_PULSES = 1 turn-on edge, latches the timer on the next rising edge and times the interval across N = op_p_pulse_count = 16 rising edges, then disables the encoder. The wide 16-edge window (vs N=2 in 1.2.4) lowers per-sample jitter; a settled loop also services far less often via the skip ladder. Two decoupled deadlines bound a sample: a short first-edge window (calc_batch_timeout_us, ~50 ms floor, scaled with SPR) declares a true stall — no encoder edge at all — quickly, while a longer measurement window (calc_meas_timeout_us ≈ 4× the 16-edge time + 200 ms) bounds completion of a moving batch. So “encoder ticks never come in” is caught in ~50 ms regardless of the coarser measurement window.

Time base. Intervals are measured in raw ticks of the ESP32-S3 internal RC slow clock (CONFIG_RTC_CLK_SRC_INT_RC, ~136 kHz nominal ≈ 7.35 µs/tick) — not the 32.768 kHz crystal. The RC oscillator drifts (±~10 % over temperature/voltage) but is continuously calibrated: the runtime stores period (µs-per-tick × 219, Q13.19 fixed-point) in RTC_CNTL_STORE1_REG, which the ULP reads via rtc_cal_period(). Conversions (shift = RTC_CLK_CAL_FRACT = 19): ticks_to_us = (ticks × period) >> 19; us_to_ticks = (us << 19) / period. The controller and its setpoint both work directly in RC ticks, so calibration is applied once (when deriving the setpoint) and clock drift cancels in the error. sample_encoder() returns raw elapsed ticks (uint32, saturating at 0xFFFFFFFF); elapsed µs are still exported as telemetry (*_elapsed_us) via ticks_to_us().

Control law — Simple-P incremental (ULP 1.3.1). Each regulated measurement computes a percent speed error and nudges the accumulated motor power by a per-pump gain, inside a 1 % deadband, saturating at the bounds and continuing (it never halts the motor on error). The accumulated power is the last DAC written (pwr = DAC_HALT − dac), seeded to start_power on the first step after a (re)start. All arithmetic is 32-bit except a single 64-bit divide for the setpoint, so no libgcc 64-bit-division helper is linked. Each update:

target_ticks = us_to_ticks( pulse_count · target_ms_per_rev · 1000 / 5184 , period )  // one udiv64

err%         = sign(meas − target) · ( |meas − target| · 100 / target_ticks )    // 32-bit hardware divu

err%         = clamp( err%, ±op_p_max_err_pct )                  // saturate + continue (default ±25)

pwr          = first_update ? start_power : (DAC_HALT − dac)       // accumulated power, clamped to [min,max]

if |err%| > 1:  pwr += err% · step_per_pct                       // 1% deadband; err>0 = too slow = +power

pwr          = clamp( pwr, op_p_min_pwr, op_p_max_pwr )           // [64, 4095]

DAC          = 4095 − pwr

The step is proportional to the percent error and shrinks as speed approaches the setpoint, so the loop converges smoothly and self-limits with no integral term — structurally eliminating the integrator-windup limit cycle of the earlier incremental-µs law (BUG 1) and the heavier positional-PID arithmetic. The measurement is clamped to 4× target_ticks before the percent divide so the magnitude can never overflow 32-bit, and target_ticks is floored at 1 to guard the divide.

Per-pump gain (host-tunable, defaults baked in). step_per_pct is the proportional gain (“kp” — power units nudged per 1 % error), set separately per pump: ac_p_step_per_pct = 7, bp_p_step_per_pct = 6. Bench tuning found AC converges crisply at 7 (holds ±1 % at the operating power) while BP — with its more compliant peristaltic load — still oscillates at 7; it is stall-free and well-damped at 6, which holds 16 s/rev tighter than the earlier 3 (hall-interval mean 15.96 s/rev, motor_stop_events = 0; bench-validated 2026-06-15 against op ULP 1.3.2). Shared knobs (all CPU-writable shared variables; defaults compiled in so the binary regulates out-of-the-box): op_p_max_err_pct = 25 (error clamp), op_p_min_pwr = 64 / op_p_max_pwr = 4095 (power bounds), op_p_pulse_count = 16 (N), op_p_skip_max = 6 / op_p_skip_good = 4 (skip ladder), op_p_start_escalate_n = 3 (no-start escalation threshold). The legacy PID gains (kp/ki/kd_scaled) remain in memory for the dosing/parking PID and host compatibility (pinned with __attribute__((used, retain)) so --gc-sections keeps them).

Adaptive skip ladder (decimation) & hall-through-skip. Once settled the loop services progressively less often, to save power and avoid fidgeting on measurement noise. After op_p_skip_good consecutive in-band (|err%| ≤ 1) samples the ladder climbs one level (cap op_p_skip_max); a panic-sized error (|err%| ≥ op_p_max_err_pct) slams it to level 0. skip_for_level(L) = 2L holds the DAC for that many wake-ticks before the next measurement (level 6 ≈ one measurement every ~1.6 s at 16 s/rev). Hall-revolution counting runs on EVERY tick — including skipped ones — so revolutions are never lost during the idle; the AC path was corrected to read the hall switch before the skip early-return, matching the BP path.

Start, no-start, and stall interventions. The regulator preserves the production start + stall machinery and adds a no-start escalation. A stall is detected fast: sample_encoder() declares one when no encoder edge arrives at all within the short first-edge window (~50 ms), decoupled from the 16-edge measurement window. The response depends on whether the motor had produced any pulse this run (has_started):

Condition

Intervention

Event bit

No-start — never produced a pulse

Spin-up KICK at DAC_START (1000 → power ≈ 3095) each tick to break stiction. If still no pulse after op_p_start_escalate_n (3) kicks, ESCALATE to full power (DAC_MAX_POWER, register 0) and hold it until the motor starts or the pump is commanded off.

MOTOR_START_FAIL_AC (bit 2) / MOTOR_START_FAIL_BP (bit 3)

Running-stall — was turning, ticks stopped

Power-cycle recovery within one tick: DAC→HALT + motor off → 200 ms pause (MOTOR_RESTART_PAUSE_US) → motor on at full power → wait one restart pulse → DAC_RESTART (full power). Repeats every tick until pulses return.

MOTOR_STOP_AC (bit 0) / MOTOR_STOP_BP (bit 1)

Both event types are written to motor_stop_events at the instant of the failure/escalation — so they are captured even if the subsequent restart succeeds — and de-duplicated per episode via recovery_signaled (cleared when a pulse returns, so a later failure logs again). The firmware reads-and-clears motor_stop_events on each hall-tick wake and logs the raw flag; the distinct bits distinguish a never-started failure from a running-stall. On the first measured pulse after any start or recovery, first_update seeds the power to start_power — the firmware's per-SPR preset captured at IDLE→RUNNING as DAC_HALT − dac_value — so the kick/restart DAC is ignored and Simple-P resumes bumplessly. stop_pump() on shutdown halts the DAC and clears the per-pump state.

Bench-validated (2026-06-10, ULP 1.3.1, target 16 s/rev):

Pump (gain)

Convergence

Settled speed

Skip level

Stall events

AC (kp = 7)

power 2850 → 470 in ~0.7 s

15.8–16.0 s (±1 %)

6 (max), held

none

BP (kp = 3)

kick → ~2400–2900 power

~16 s mean, looser ripple

climbs to 6

none

AC converges crisply and holds the setpoint to the measurement-noise floor, with revolutions counted correctly through the skip-level-6 idle. BP at kp = 7 oscillated and produced one running-stall that recovered cleanly (validating the recovery + event path); at kp = 3 it is stall-free. The residual BP ripple is the peristaltic roller pulsation (physical), not gain-removable — consistent with the earlier PID Ki-sweep finding.

DAC output. Written to DAC7574 (0x4C) via I²C on GPIO1/GPIO2 (ULP-accessible). DAC = 4095 − power (active-low: higher power = more drive). DAC_START = 1000 → power ≈ 3095 (aggressive startup kick).

Legacy PID retained. The positional integer-tick PID (run_pid(), from ULP 1.2.4) and its gains (kp/ki/kd_scaled, bench-tuned 300/120000/0) remain in the operational binary for the two non-SPR modes that still need them: BP precision dosing (BP_DOSING, counts encoder edges to a target) and hall parking/calibration (BP_PARKING). Steady-state speed regulation no longer uses it.

Hall switch edge detection and CPU wake. During BP_RUNNING, the ULP reads blood pump hall switch (GPIO21, active LOW, DRV5032AJDMRR) every tick. On 1→0 transition: increments bp_hall_revolutions (uint32) and fires wakeup_main_cpu() (RTC_CNTL_SW_CPU_INT). Edge tracker reset to 1 on BP_IDLE→BP_RUNNING. 32-bit counter: ~2 315 years to overflow at 17 s/rev.

Shared variables (.bss.priv section, addresses auto-generated to variables.py):

Variable

Type

Direction

Description

bp_hall_revolutions

uint32

ULP → CPU

Cumulative revolution count

main_cpu_state

uint32

CPU → ULP

Main CPU state machine state

bp_hall_switch

uint32

ULP → CPU

Hall level (0=active, 1=open)

bp_elapsed_us

uint32

ULP → CPU

Last encoder window µs

bp_target_ms

uint32

CPU → ULP

Target SPR in ms

motor_stop_events

uint32

ULP → CPU

Stop event bitmask (read-and-clear)

bp_state

uint32

ULP → CPU

IDLE / RUNNING / PARKING

Main core interface. blood_pump.spr = N, blood_pump.enable(True/False), blood_pump.revolutions. MicroPython class ULPBloodPump in pump_regulator.py.

Source: esp32/pump_regulator/main/ulp/main.c, fluispotter/variables.py, fluispotter/pump_regulator.py. Build: esp32/pump_regulator/ULP_BUILD.md.

Cooperative AC service during BP dose/coast/park (ULP 1.12.0, SWR-271): BP_DOSING (~10 s), BP_DOSE_COAST (5 s) and BP_PARKING (≤30 s) run as blocking loops that previously starved the AC pump entirely — an AC hall pass falling inside such a window was lost outright (false “did not advance” warnings, ~3/h in the 8 h field run) and the AC ran open-loop on its held DAC. The ULP now (a) edge-polls the AC hall inside every blocking BP spin with semantics identical to ac_tick, and (b) runs a throttled (every 16th window, ~0.8 s), hard-capped (150 ms), 8-interval AC Simple-P regulation slice between BP drive windows in DOSING and PARKING; the slice’s combined poll keeps counting BP encoder edges into the dose/park ledger so dose volume accounting stays EXACT, a 160-edge tail guard keeps target-crossing out of any slice, and the park slice watches the BP hall so home is never overshot. Bench proof: ac_busy_slices 0→18 with the AC DAC live through one real dose+park, dose exact (3070/3067, +3 = coast edges). Field: 13 samples, ZERO hall warnings.

Pump run-monitor (fw 0.1.110, SWR-270): the CPU-side stall monitor watches the ULP hall-revolution counters (AC/BP_HALL_REVOLUTIONS, RTC domain) fail to advance within 2.5×SPR — never a CPU GPIO edge, which a coincident BP IRQ burst could starve into false warnings (48 on 2026-07-13).

RTC-pad register-clock defense (#48 — “motors run wild after USB-attached reboots”, root-caused + fixed 2026-07-22/23): SENS.sar_peri_clk_gate_conf.iomux_clk_en (0x60008904 bit 31) is the register clock shared by ALL 22 RTC pads. ESP-IDF’s rtcio_ll_function_select() clears it globally whenever the main CPU routes ANY pin to digital (machine.Pin() → rtc_gpio_deinit — valve toggles, ADC reads). With the clock gated, every ULP rtcio write is silently lost — encoder blind, enables dead — while all pad-config registers READ sane; the designed kick escalation then drives to full power on zero tacho (the “wild motors”), and the next lightsleep’s ext1_wakeup_prepare re-enables the clock (the self-healing). Convicted by ULP-side capture at the failed kick (clock bit read 0). Fix: the ULP re-asserts the clock at the top of every 1 ms tick (DEFEND_CLOCK(): SET_PERI_REG_MASK, RMW) — pump-pad ownership enforced at silicon level per the owner principle “pumps entirely under ULP control”; a CPU gate now costs at most one tick (residual: a gate inside one kick/drive window costs one benign single-step retry). Tick-top ONLY: placing the call inside the deep call chain forces +16 B stack frames per formerly-leaf level, invisible to static headroom accounting — the per-poll-iteration variant canary-tripped 5/5 recovery boots and the window-entry variant intermittently (a canary-tripped ULP keeps its core loop advancing but silently fails pump starts — issue #49). Proven on the TestOp dev track (a fourth binary, pump_testop.bin, MAJOR 9 rendered dev.MINOR.PATCH: operational code + instrumentation, iterated while pump_regulator.bin stays frozen at last-known-good): dev.0.3 3/3 clean recovery trials vs an 11/12-blind baseline; dev.0.6 survived 7+ live WDT-recovery boots and two full sessions on 2026-07-23. Roll-in to operational ULP 1.16.0 is code-complete, gated on the stack-headroom budget decision (552 B built vs the 768 B build-gate line; the empirical canary trip line is ≤480 B). GitHub #48; SWR-114 (DOC 1.88) documents the dev track.

5.6 Hall-Wake Session Architecture

The session loop uses hall-tick-driven Light Sleep: the main CPU sleeps indefinitely, waking when the ULP fires wakeup_main_cpu(). The ULP wakes the CPU on each blood-pump hall revolution, on each AC-pump hall revolution (ac_hall_revolutions), AND on a motor stop/stall event (motor_stop_events) — not BP-only as earlier revisions implied. CPU draws ~2 mA sleeping, wakes for ~1–5 ms housekeeping. (During DOSING the loop uses a plain asyncio 50 ms tick, not Light Sleep.)

Wake handler (_on_hall_tick)

On each wake:

1. Error check — read motor_stop_events; log warning if non-zero, clear flag.

2. Deque update — read bp_hall_revolutions, compute new revolutions, append external RTC timestamp to _blood_pump_sample_history (deque length 7). Proves blood from catheter tip at scheduled time has traveled 7 pump revolutions to dosing point.

3. Counter update — increment blood_pump_revolutions for telemetry.

4. External RTC verification — every 10 revolutions, log MCP79410 RTC read confirmation (IC6, I²C 0x6F, 32.768 kHz crystal Y2, ±2 ppm).

Beyond these BP steps, _on_hall_tick also reads + accumulates the AC-pump revolution counter (ac_pump_revolutions) with AC SPR diagnostics, and logs a per-wake “Hall tick” entry.

Session loop pseudocode

sub = IDLE

while True:

_on_hall_tick()                              # ALWAYS serviced

if sub != DOSING: sync_on_new_revolution()

if sub == IDLE:

if remaining ≤ 5 min: sub = TRIGGER   # wake up

elif sub == TRIGGER:

if deque_confirmed(): sub = DOSING       # 7-rev deque OK

dose_ctx = _dosing_init()

elif sub == DOSING:

if _dosing_tick(dose_ctx) == DONE:

if more_doses: dose_ctx = _dosing_init()

elif more_samples: sub = IDLE

else: return FINISH

await sleep(500 ms if sub != DOSING else 50 ms)

The deque stores RTC timestamps of hall-switch events (one per revolution). When all 7 entries post-date the scheduled sample time, the pump has completed at least 7 full revolutions since that time. At 16 s/rev this is ~112 seconds of continuous pumping, confirming fresh catheter blood is at the nozzle.

Timing

Wake frequency (16 s/rev)

~5 400 wakes/day

Wake duration

~1–5 ms

CPU active duty

< 0.03 %

Deque confirmation delay

~112 s (7 × 16 s) after scheduled time

Light Sleep current

~2 mA

Active current (240 MHz)

~50 mA (§4.3 upper bound; was 80 MHz / ~31 mA)

5.6.1 Light-Sleep State Map (Power Management — SWR-243..250)

Light Sleep is gated by a phase rule: the firmware enters machine.lightsleep() only once the device has reached the SESSION state (state number ≥ _STATE_SESSION = 10). Every pre-session state performs its waits as plain asyncio sleeps — keeping the radio and the event loop alive so an operator/hub command is serviced within the ~1 s poll interval, and so no sleeper can run while an actuator is being driven. From SESSION onward the organizing structure is the blood-pump flush cycle (SWR-243): between samples the BP runs 2 minutes at 16 SPR, then parks at the sealed dead-point and rests 8 minutes, repeating — and it CONTINUES in FINISH so the line stays flushed after the last sample (SWR-248). The CPU light-sleeps in every phase except active dosing and fault handling (SWR-245); when the pump is turning the ULP wakes it on each hall edge just to update the length-7 deque (SWR-247), then it sleeps again. The map below is the fw 0.1.86 behaviour.

State (#)

Light Sleep

Sleep cadence

Wakes on

Wi-Fi held

INITIAL (1)

No

recovery I/O only

sync / file I/O completion

ephemeral (recovery)

AWAIT_CALIBRATION (2)

No

~1 s poll (asyncio)

poll timer; acquire; command

Yes

CALIBRATION (4)

No

50 ms tick (asyncio)

valve switches / tape timeout / parking test done. Entry is gated by the ULP health verify (SWR-250): stack-overflow canary clear + core loop advancing, else abort with a POWER-CYCLE error.

No

AWAIT_ANTICOAGULANT_BAG (5)

No

~1 s poll (asyncio)

command

Yes

ANTICOAGULANT_BAG (6)

No

~1 s poll (asyncio)

command

Yes

AWAIT_CATHETER (7)

No

~1 s poll (asyncio)

command; AC-cycle 1 s tick

Yes

CATHETER (8)

No

500 ms detect / ~1 s poll (asyncio)

blood threshold or command

Yes

AWAIT_SESSION (9)

No

~1 s poll (asyncio); flush cycle runs on plain asyncio pre-session

START_SESSION command

Yes

SESSION · IDLE — flush RUN (10/0, 2 min)

Yes (flush cycle)

lightsleep between hall ticks (≤60 s backstop; RUN window RTC-timed)

ULP wake per BP hall edge (~16 s) → deque update (SWR-247); motor stall / failed start (immediate); RTC backstop

end-of-run sync only (SWR-246)

SESSION · IDLE — flush REST (10/0, 8 min)

Yes (flush cycle)

lightsleep(≤60 s backstop, clamped to the preload boundary T−5 min); motor parked + OFF; ~20 mA bench-measured

RTC backstop (no hall ticks); stall/fail if any; ends early when the next sample is <5 min out

down (association dropped)

SESSION · TRIGGER — preload/loading (10/1)

Yes (session loop; flush cycle SUSPENDED)

lightsleep between hall ticks (60 s backstop); BP runs continuously at 16 SPR (SWR-244)

ULP wake per hall edge — appends one RTC timestamp per revolution to the length-7 deque; dose triggers when 7 revolutions post-date the sample time (all state changes are hall-edge-driven)

down; the pre-dose sync re-associates

SESSION · DOSING (10/2)

No

50 ms tick (asyncio)

dosing state-machine tick; post-dose sync (SWR-246)

re-associated for pre/post-dose sync

FINISH (12)

Yes (flush cycle continues — SWR-248)

flush RUN/REST as in SESSION IDLE; the SWR-241 heartbeat loop stays on plain asyncio and rides along frozen

as flush RUN/REST; heartbeat re-syncs + polls commands every ~10 min (RTC deadline)

per heartbeat + flush-cycle syncs

Single-owner backstop light-sleep

At most one task sleeps the CPU at any time, via one helper idle_lightsleep(backstop) → machine.lightsleep(LIGHTSLEEP_BACKSTOP_MS) (default 60 s, board_config-tunable, reset each nap). The two sleepers are mutually exclusive by construction: the flush cycle owns the nap through its RUN and REST phases (SESSION IDLE and FINISH), and the session loop owns it during preload/loading (TRIGGER) — which only exists while the flush cycle is suspended. There is no scheduler, no election, no registration — the earlier timed-election coordinator was removed entirely after it hung the CALIBRATION state (see the regression note). The ULP wakes the CPU EARLY on any hall revolution, motor stall, or failed start (wakeup_main_cpu → RTC_CNTL_SW_CPU_INT); the backstop is the RTC-timer liveness bound if no event fires. After every nap the sleeper yields _LIGHTSLEEP_SETTLE_MS (50 ms) so any frozen concurrent task runs one iteration before it sleeps again. A fault report in flight, or a pending start-FAIL confirmation, sets inhibit_lightsleep_until_ms and the sleeper degrades to short asyncio naps so those timers are not stretched.

Other tasks never sleep the CPU

Telemetry, the AC pump cycle, and the session / FINISH loops (outside TRIGGER) use plain asyncio.sleep_ms and never call machine.lightsleep — they simply ride along frozen during a nap and resume on wake. The AC pump keeps regulating on the ULP throughout every nap (bench-verified: AC holds SPR through the flush cycle). DOSING (SESSION sub-state 2) stays awake on a 50 ms asyncio tick — the CPU never sleeps while actively dosing.

Batched hub sync (SWR-246)

In-session hub sync is batched, not per revolution: the flush cycle syncs at each RUN boundary (_synchronize_from_blood_pump_cycle opens a fresh station, so an association dropped by the REST nap is re-established), the session loop syncs before and after each dose (_sync re-associates if needed), and faults push immediately through their own bounded station (SWR-236/249). The hub therefore updates roughly every 2-minute RUN plus at each dose and on errors; events that cannot be delivered at a sync point stay committed to flash and ship at the next one. The radio is fully down through the 8-minute REST — the association is deliberately not preserved across naps.

Real-time correctness across a nap

Because machine.lightsleep freezes time.ticks_ms (but not the external RTC), every deadline that must survive a nap is measured on the RTC: the sample schedule (_next_sample_ms), the FINISH 10-minute heartbeat, and the blood-pump idle window all use _rtc_time_ms(). The length-7 loading deque is count-based — _on_hall_tick appends one RTC timestamp per revolution and TRIGGER→DOSING fires when 7 revolutions post-date the sample time — so it is robust to how coarsely the CPU wakes.

ULP faults wake for logging

The ULP raises a CPU wake (wakeup_main_cpu()) not only on each hall revolution but also on a motor stall (MOTOR_STOP_AC/MOTOR_STOP_BP) and a failed start (MOTOR_START_FAIL_AC/MOTOR_START_FAIL_BP), ORed into motor_stop_events. On every wake the owner's first housekeeping step (_on_hall_tick → _drain_motor_stop_events) reads-and-clears the bitmask, decodes it, and writes a warning/log event with both pumps' drive + SPR (SWR-236). A fault is therefore logged promptly — near-instant if the fault interrupt cuts the nap short (OI-12), and at most one backstop (60 s) otherwise. Faults are never lost to sleep. A wake-side log marker (“Woke from lightsleep”) follows every nap, so a CPU wedged inside machine.lightsleep shows an “Entering…” with no matching “Woke…”.

Terminology

asyncio.sleep_ms(n) is the stock MicroPython cooperative sleep: it yields the current coroutine for n ms while the event loop keeps running other tasks — the CPU stays fully awake, no power saving. Used everywhere except the owner's nap: DOSING, the worker tasks, and every pre-session poll.

idle_lightsleep(backstop) is the firmware's sleeper helper: it powers the CPU down with machine.lightsleep(backstop) so the whole device (CPU + event loop) stops until the ULP wakes it or the backstop timer elapses. Its only callers are the flush cycle (RUN + REST) and the session loop's TRIGGER branch.

“Untimed + backstop”. The nap is conceptually untimed — it is meant to end when a real event (hall / stall / fail) fires the ULP wake — but it is bounded by the backstop (60 s) as liveness insurance so a missed wake can never hang the device: it wakes at least once per backstop, re-checks everything on the RTC, and sleeps again. This restores the §5.6 model (untimed, ULP-woken) with a safety timer, replacing the rejected timed-election coordinator.

Regression note. An earlier fw 0.1.80 build used a multi-sleeper timed election and routed pre-session telemetry/command-poll waits through it, which made the election run in the CALIBRATION state where it had never run before, stalled the calibration loop, and left the tape reel energized (it ran continuously). That model was rejected. Option A removes the election entirely and only sleeps the CPU from the main loop at state ≥ SESSION, so no pre-session phase can sleep or run a shared coordinator — the invariant the incident violated.

Wake-source / power reconciliation (OI-12 — largely confirmed). Bench 2026-07-07: the flush REST window measures ~20 mA with the CPU light-sleeping while the ULP keeps regulating the AC pump; the ULP wake from machine.lightsleep is confirmed (hall-tick deque updates and immediate stall/fail reporting through naps). Remaining: PPK2 profile of the RUN phase and the preload window under fw 0.1.86 (both now sleep between hall ticks), and a full-session energy figure. See §8.1 OI-12.

5.6.2 Session Recovery (Power-Loss Resume)

A power loss during a session MUST NOT lose the session. The device persists its session state to flash and, on every boot, resumes where it left off — guarded (fw 0.1.92/0.1.93, SWR-257/258) so that a resume can never run blind on a dead pump controller or pump blood without anticoagulant.

Persisted state

/state.dat — device state number, current sample index, session start time, blood-pump and AC-pump revolution counters, and the acquired hub’s SSID/BSSID. Written on every persisted-state change (session start, sample advance, FINISH entry). /time-table.dat — the sample schedule, written when settings are received from the hub. Deleting both files (the Control app’s Session Reset, SWR-317) is the sanctioned way to abandon a session: the next boot then initializes clean into AWAIT_CALIBRATION.

Boot algorithm (_state_initial)

1. Valve close (2 s cap, best-effort) — a crashed session may have left it open.
2. Restore /time-table.dat and /state.dat. Either missing or unreadable → clean initialization → AWAIT_CALIBRATION.
3. If the saved state is SESSION or FINISH: best-effort hub reconnect — ONE attempt to associate to the saved SSID/BSSID, NTP (accepted only inside the SWR-256 plausibility window), event upload and command fetch. Any failure is logged and the session resumes OFFLINE (it syncs at the next planned connect) — recovery is never blocked by the network.
4. ULP health verify (SWR-250/257): stack-overflow canary clear AND the core loop ADVANCING. On failure an error naming the required POWER-CYCLE is reported, the session is NOT resumed, and the device falls back to AWAIT_CALIBRATION — hub-visible, commandable, pumps off. This is underpinned at a lower level by the ULP start gate itself (SWR-257): the loader zeroes the INITIALIZED counter BEFORE starting the core and requires it to advance (≥ 2 invocations) — the random RTC-memory noise a power cut leaves behind cannot fake a started ULP (the pre-0.1.92 nonzero-read check could be, and was, fooled). One reload retry, then a hard failure.
5. Restore counters and schedule position (sample index + session start time). Resume does not consume a sample.
6. Start the AC pump cycle, followed by one explicit scheduler yield so its enable executes before the state machine moves (SWR-258). The session-entry path below runs synchronously into the blocking machine.lightsleep; without the yield a freshly-created task would not run until the first wake — the pre-0.1.93 defect that let the blood pump run 20–60 s without anticoagulant on resume.
7. Enter SESSION (or FINISH if the saved state was FINISH — the flush cycle continues per SWR-248; the session loop is not entered).

Session entry on resume

If the next sample is inside the 5-minute preload window (after a mid-session power cut it is typically overdue), the session takes ownership of the blood pump BEFORE the entry sync (SWR-244) — a slow hub cannot delay the preload. Independently, EVERY blood-pump start enforces the AC-first invariant (SWR-258) at the ULP register level: in any state ≥ SESSION, if the AC pump is not running it is started first (same SPR/preset/enable sequence the AC cycle uses) and a warning event is reported — on a healthy resume this guard stays silent; its warning appearing in the hub log means the primary ordering failed and should be investigated. The guard is inert below SESSION, so CALIBRATION and the priming states keep their designed dry/manual AC behaviour.

Operational note (bench, 2026-07-08). A session that has been recovered many times carries its ORIGINAL start time; once wall-clock time passes multiple schedule slots, every elapsed sample is overdue at once and the device processes the backlog back-to-back — rapid load/dose/park cycling that presents as continuous short tape/valve moves. This is the schedule catching up, not a fault; for bench work prefer a Session Reset and a fresh session over repeated recoveries of a stale one.

5.6.3 Hall-Deque Integrity — Registration Proof & Crossing-Time Stamps (fw 0.1.130/0.1.131)

What the deque is: the ULP timestamps every BP hall crossing (in EVERY state, SWR-289) into a 9-slot ring; the CPU consumes the ring into a 7-deep deque of crossing times. sampled_at semantics: the deque stamp marks when the blood now at the tape valve entered the catheter at the tip; it then transits the line for 7 BP revolutions before reaching the tape (the §5.7 load gate: dose only when the oldest of the last 7 crossings is at/after the scheduled time AND has advanced past the previous sample’s anchor — so backlog drains at one sample per revolution plus the ~50 s dose choreography, and every spot carries post-schedule blood).

Registration is proven complete (owner-requested audit, 2026-07-23): the overlay tunable DEQUE_AUDIT (board_overrides.json, default 0) makes the consume path emit one event per deque entry plus a per-batch completeness line (ring head vs ULP revolution counter vs CPU consumed total vs deque length) and an always-on warning when the 9-slot ring overflows. Live verdict across two sessions and 7 reboots: 120/120 and 70/70 crossings accounted, every contiguous segment exact, zero ring overflows — no hall tick is ever lost by the ring→deque→counter chain, including across WDT-reset recovery.

Crossing-time stamps are trustworthy since fw 0.1.131 (issue #51 — the audit caught three corruption modes live and all three are closed): (1) wrap clamp — the ULP publishes now_ticks once per tick and the publish goes stale through its busy windows (dose/park up to 30 s), so a crossing could read NEWER than “now” and the unsigned subtraction wrapped to a stamp 2³²/rate ms in the past (observed: exactly 105,703,256 ms = 2³²/40.617); a half-range test now stamps such entries at poll time (error bounded by publish staleness, seconds not hours). (2) rate hygiene — the tick-rate self-calibration runs AFTER the batch, is skipped (and its reference reset) when the batch clamped, and accepts only 110–170 ticks/ms (RC_SLOW nominal 136 ±25%; the stale-pair regime produced 40.6/69.2 undershoots and a 201.5 overshoot echo). (3) seed — the rate initialises at the 136 ticks/ms nominal so the first batch after boot/recovery back-dates within clock tolerance instead of stamping “now”. Acceptance (2026-07-23, 6 segments / 5 reboots): 0 negative deltas, 0 far-past stamps, all rates in-band, 5 wraps fired and all 5 neutralized. Residual accuracy: stamps are good to ≤ publish staleness (seconds) — adequate for the minutes-scale decisions built on them.

Reboot cost: recovery clears the CPU-side deque (stamps are RAM), so the first sample after every reboot pays a full 7-revolution refill (~96 s at 16 SPR) before dosing — deliberately conservative (re-establishes the post-schedule guarantee). Between reboots, consecutive backlog samples dose with ZERO fresh-revolution wait (proven live: loaded 7/7 in the same second, 8 s after the previous sample completed).

5.7 Blood Sampling Process

This section describes the complete blood sampling sequence executed by the main CPU firmware during a session. The process is driven by a pre-configured time table (list of sample times and dose counts received from the hub server) and coordinates the ULP pump controller, valve motor, and tape reel motor.

5.7.1 Constants

Parameter

Value

Notes

Encoder pulses per output revolution

5 184

324 (gear ratio) × 16 (encoder edges; exact 323.628 × 16 = 5 178). The prior “326” was a typo — 326×16 = 5 216 ≠ 5 184.

Blood pump SPR (target)

16 s/rev

BLOOD_PUMP_SECONDS_PER_REVOLUTION

Prepare degrees

50°

_BLOOD_PUMP_SAMPLE_PREPARE_DEGREES

Prepare encoder pulses

720

round(5184 × 50 / 360)

Collect degrees

213°

_BLOOD_PUMP_SAMPLE_COLLECT_DEGREES

Collect encoder pulses

3 067

round(5184 × 213 / 360)

Deque confirmation length

7 revolutions

_BLOOD_PUMP_SAMPLE_HISTORY_LENGTH

Valve drive timeout

10 000 ms

_VALVE_DRIVE_TIMEOUT_MS

Tape reel power

100 (of 4 095)

Open-loop; no encoder feedback

Tape reel timeout

5 000 ms ⚠

Per drive operation (smear + advance) — _TAPE_REEL_TIMEOUT_MS = 5000 in code (was documented as 20 000 ms; under review — is 5 s enough to smear/advance the tape? §8.1).

Session wake-up threshold

5 min (300 000 ms)

Sub-state IDLE → TRIGGER within session loop

Flush cycle — active

120 000 ms (2 min)

_BLOOD_PUMP_CYCLE_RUNNING_MS

Flush cycle — idle

480 000 ms (8 min)

_BLOOD_PUMP_CYCLE_IDLE_MS

5.7.2 Timeout Formula

All blood pump movement timeouts use a single formula derived from the target SPR and a 30× safety margin:

degree_timeout_ms = (SPR × 30 × 1000) / 360

= (16 × 30 × 1000) / 360

= 1 333.33 ms per degree

timeout(degrees) = int(1 333.33 × degrees)

Movement

Degrees

Timeout

Prepare (ready position)

50°

66 666 ms (~1.1 min)

Collect (blood onto tape)

213°

284 000 ms (~4.7 min)

Park (full rotation)

360°

480 000 ms (8 min)

5.7.3 Flush Cycle (Pre-Session Background Task)

Before the SESSION state begins sampling, the blood pump runs a repeating 2+8 flush cycle to keep the catheter clear. The cycle is managed by _run_blood_pump_cycle() and reported as three sub-states: ACTIVE, IDLE, SUSPENDED.

1. ACTIVE (2 min) — ULP regulates the blood pump at 16 s/rev. Simple-P regulation runs continuously (see §5.5).

2. Park — At the end of the active phase, the pump is parked at the hall home position (flush-cycle wall-clock backstop 34 s — _park_deadline in _run_blood_pump_cycle; the prior “480 s” was incorrect for this flush-cycle park, which is distinct from the dose-sequence park timeout).

3. IDLE (8 min) — Pump disabled. Hub sync performed at start and end of idle window.

4. Loop — Returns to step 1.

When SESSION enters the sampling loop, the flush cycle task is cancelled and its state set to SUSPENDED. The ULP runs the pump continuously from that point on.

5.7.4 Unified Session Loop (STATE 10: SESSION)

On entry, the firmware:

1. Cancels the flush cycle task (if active).

2. Enables the blood pump via ULP at 16 s/rev. The pump now runs continuously — no 2+8 cycling.

3. Registers the hall-switch interrupt handler (_on_hall_tick).

4. Enters a single polling loop with three sub-states: IDLE (0), TRIGGER (1), DOSING (2).

Each iteration (500 ms cadence in IDLE/TRIGGER, 50 ms in DOSING):

• Always calls _on_hall_tick() — reads ULP revolution counter, appends RTC timestamps to the deque.

• Outside DOSING: syncs with hub on every new hall revolution.

• Sub-state dispatch:

IDLE (sub 0) — Waits for scheduled sample time. When remaining_ms ≤ 5 minutes, transitions to TRIGGER.

TRIGGER (sub 1) — Waits for 7-revolution deque confirmation. Once the oldest deque entry post-dates the scheduled time, 7 full revolutions of fresh catheter blood have reached the nozzle (~112 s at 16 s/rev). Transitions to DOSING.

DOSING (sub 2) — Executes _dosing_tick(ctx) (see §5.7.5). When all doses for the current sample complete, advances the time table. If more samples remain, returns to IDLE. Otherwise transitions to FINISH (state 12).

5.7.5 Dose Collection — _dosing_init() + _dosing_tick(ctx)

Each dose is initialised by _dosing_init() which returns a context dict, then driven by _dosing_tick(ctx) called once per 50 ms loop iteration. The tick function returns _DOSE_CONTINUE or _DOSE_DONE.

5.7.6 Dose Step Table

Each dose is executed by a tick-based state machine (50 ms tick). Every sub-action is decomposed into an instant start (GPIO/mem32 write), a synchronous poll (GPIO/mem32 read), and an instant stop — no dedicated busy-wait per actuator. Independent actuators (tape reel, valve) run concurrently where safe.

The prepare and collect steps use ULP-managed dosing (bp_dosing / bp_dose_target) instead of CPU-side pulse counting, offloading encoder counting entirely to the ULP co-processor. The CPU polls a single flag per tick.

Step

State

Action

Start

Poll (each tick)

Timeout

0

PARK_1

Park — Hall home position

bp_parking = 1

bp_parking == 0

480 000 ms

1

ENSURE_CLOSED

Verify valve closed

direction(CLOSE); power(P)

close_switch == 0

10 000 ms

2

PREPARE

Dose 720 pulses (50°)

bp_dose_target=720; bp_dosing=1

bp_dosing == 0

66 666 ms

3

OPEN_VALVE

Open valve

direction(OPEN); power(P)

open_switch == 0

10 000 ms

4

COLLECT

Dose 3 067 pulses (213°)

bp_dose_target=3067; bp_dosing=1

bp_dosing == 0

284 000 ms

5

SMEAR_AND_CLOSE

Concurrent: tape smear + valve close

tape.power(100); valve.direction(CLOSE); valve.power(P)

close_switch & elapsed

valve 10 s / tape 5 s

→ On valve closed: enable BP immediately

6

ADVANCE_TAPE

Advance tape (BP already running)

tape.power(100)

elapsed ≥ 5 s

5 000 ms

7

PARK_2

Park (preempts BP_RUNNING)

bp_parking = 1

bp_parking == 0

480 000 ms

Concurrency gain: the valve close (~10 s) and tape smear (~5 s, _TAPE_REEL_TIMEOUT_MS) run concurrently in step 5, and the BP restarts the instant the close-switch trips rather than waiting for the tape. (With the current 5 s smear the valve close is the longer leg; the gain is largest if the tape window is later raised.)

Worst-case single dose duration: 480 + 66.7 + 10 + 284 + max(5, 10) + 5 + 480 = 1 335.7 s ≈ 22.3 min (using the code's 5 s tape windows).

5.7.7 ULP Integration — Park and Dose

Parking (bp_parking / bp_parked):

1. CPU writes bp_parking = 1 in RTC shared memory.

2. ULP enters BP_PARKING state. Motor drives toward hall switch at PID-regulated speed.

3. On hall activation (GPIO LOW), ULP counts bp_park_offset_ticks additional encoder edges to centre the rotor in the hall-active zone.

4. ULP clears bp_parking and sets bp_parked = 1 (success) or bp_parked = 0 (failure). Failure is detected by encoder rotation, not wall-clock time: the ULP accumulates encoder ticks into bp_park_enc_total and declares a timeout once it has turned a full revolution (PARK_ENC_TICK_LIMIT = 5200 ticks ≈ 1 rev) without the hall switch firing. A 30 s hard wall-clock backstop (PARK_SAFETY_TIMEOUT_US) guards against a stalled motor that produces no ticks, and the post-hall offset coast uses a 5 s backstop (PARK_OFFSET_SAFETY_US). This replaces the previous fixed 25 s PARK_TIMEOUT_US, which could report success even when the motor never started. bp_park_enc_total is exposed to the CPU so calibration can distinguish “motor did not start” (0 ticks) from “hall switch not detected” (≥ 5200 ticks, motor confirmed turning).

5. State machine polls bp_parking each 50 ms tick until cleared.

Dosing (bp_dosing / bp_dose_target / bp_dose_actual):

1. CPU writes target encoder pulses to bp_dose_target and sets bp_dosing = 1.

2. ULP enters BP_DOSING state. Counts encoder edges at PID-regulated speed.

3. When target reached, ULP transitions to BP_DOSE_COAST (1 s capture window for overshoot).

4. ULP clears bp_dosing, writes actual count (incl. coast overshoot) to bp_dose_actual, returns to BP_IDLE.

5. State machine polls bp_dosing each 50 ms tick until cleared.

5.7.8 State Transition Diagram

SESSION (state 10) — unified loop with sub-states

└ pump runs continuously at 16 s/rev via ULP

└ _on_hall_tick() serviced EVERY iteration

└ syncs with hub on every new hall revolution (outside DOSING)

│

└ sub=IDLE (0): polls 500 ms, waits for schedule

│   └─ remaining ≤ 5 min → sub=TRIGGER

│

└ sub=TRIGGER (1): polls 500 ms, waits for deque

│   └─ 7-rev deque confirmation → sub=DOSING

│

└ sub=DOSING (2): _dosing_tick(ctx) at 50 ms cadence

│   └ for each dose (1..N):

│   │   0: PARK_1        → park at hall home

│   │   1: ENSURE_CLOSED → verify valve closed

│   │   2: PREPARE       → dose 720 pulses (50°)

│   │   3: OPEN_VALVE    → open to limit switch

│   │   4: COLLECT       → dose 3 067 pulses (213°)

│   │   5: SMEAR_AND_CLOSE → tape + valve concurrent

│   │                         valve closed → enable BP

│   │   6: ADVANCE_TAPE  → next slot (BP running)

│   │   7: PARK_2        → park at hall home

│   └ all doses done → advance time table

│       └─ more samples → sub=IDLE (loop)

│       └─ all complete → FINISH (state 12)

5.7.9 Blood Pump Motor-Start Paths

Seven distinct code paths start the blood pump motor. They fall into two categories: state-machine paths that run the regulation loop with stall recovery (Simple-P for the running state since ULP 1.3.1; the legacy positional PID for dosing/parking — see §5.5), and test/diagnostic functions that drive the motor at a fixed or stepped DAC level by design. Every path that depends on rotation now verifies the motor actually turned via encoder ticks (see §5.7.7).

A. Paths that run the regulation loop (regulate_pump() or its inline PID)

Path

Regulation law

Stall recovery

Notes

BP_RUNNING

Simple-P (regulate_pump())

Yes (+ no-start escalation)

Normal continuous operation at target SPR (see §5.5)

BP_DOSING

Inline PID (same logic)

Via DAC nudge

Counts dose encoder ticks to bp_dose_target

BP_PARKING

Inline PID (same logic)

Via DAC nudge

Drives toward hall switch; verified by encoder ticks

B. Test / diagnostic functions with their own motor start (fixed/stepped DAC, no PID)

Path

DAC drive

Stall recovery

Purpose

test_determine_parking

PARK_CAL_DAC (1095)

None — encoder-tick verified

Blocking loop measuring the hall bracket; counts ticks (5200 limit) so a non-starting motor is detected rather than silently timing out

test_resistance_map

Variable (stepped)

None

Steps through DAC levels measuring resistance per segment; not PID by design

test_fixed_power

Constant

None

Holds a constant DAC level; no PID by design

test_verify_encoder

DAC_MAX_POWER (0)

None

Full power, counting encoder ticks per revolution for fidelity check

The state-machine paths (A) already mirror regulate_pump(), so they were not consolidated into a single function. The key robustness change is that every rotation-dependent path now verifies the motor physically turned via encoder ticks rather than trusting a wall-clock timeout.

5.8 Fatal Error Messages

When a fatal timeout occurs during device operation, the firmware logs a human-readable error message to the event store (synced to hub) and writes it to fatal-error.txt. The device enters an error blink state (LED toggling at 500 ms) until reset via reed switch.

Error code

Stage

Message

ensure_closed

Sample: valve close before dose

Valve position detection failed. Possible mechanical break of valve arm axle connector. Return unit for replacement.

valve_open

Sample: valve open for blood

Valve position detection failed. Possible mechanical break of valve arm axle connector. Return unit for replacement.

valve_close_smear

Sample: valve close during smear

Valve position detection failed. Possible mechanical break of valve arm axle connector. Return unit for replacement.

park_1

Sample: park BP before dose

Blood pump failed to reach parking position. Hall effect sensor not detected within expected rotation. Check pump motor, encoder, and hall sensor connections. Or return unit for replacement.

prepare

Sample: dose prep (50°)

Blood pump failed to complete dose preparation (50 deg). Pump motor or encoder malfunction. Check motor drive and encoder signal. Or return unit for replacement.

collect

Sample: blood collection (213°)

Blood pump failed to complete blood collection (213 deg). Pump motor or encoder malfunction. Check motor drive and encoder signal. Or return unit for replacement.

park_2

Sample: park BP after dose

Blood pump failed to reach parking position. Hall effect sensor not detected within expected rotation. Check pump motor, encoder, and hall sensor connections. Or return unit for replacement.

valve_close_init

Startup: valve close (_state_initial)

Valve position detection failed. Possible mechanical break of valve arm axle connector. Return unit for replacement.

valve_open_ac_bag

Setup: valve open (_state_await_anticoagulant_bag)

Valve position detection failed. Possible mechanical break of valve arm axle connector. Return unit for replacement.

valve_close_catheter

Setup: valve close (_state_await_catheter)

Valve position detection failed. Possible mechanical break of valve arm axle connector. Return unit for replacement.

The same _VALVE_TIMEOUT_MSG also appears during calibration — there it is non-fatal (calibration continues to the next step) — and during the three pre-session state-transition codes added above (valve_close_init, valve_open_ac_bag, valve_close_catheter). In the state transitions the timeout IS fatal: it is raised, caught by run_safe → written to fatal-error.txt + the diagnostics CSV, and the device enters the error blink until reset — identical to the in-sample codes. (Only the calibration appearance is non-fatal; earlier wording ambiguously grouped the two.)

6. Synchronize Service — Raspberry Pi Hub Server

6.1 Overview & Architecture

The synchronize service is a Python web application running on a Raspberry Pi that acts as the central hub for all Fluispotter devices. It provides a REST API for device telemetry upload, event logging, settings distribution, and file transfer, plus a web UI for operators to manage devices and monitor sessions.

Language

Python 3 (async)

Framework

aiohttp (async HTTP server)

Database

SQLite via aiosqlite (async)

Validation

marshmallow (request/response schemas)

Templating

Jinja2 (web UI pages)

Port

9090 (env var HTTP_PORT)

Data directory

/data/synchronize-service/ (env var PATH_BASE_DIRECTORY)

Source location

lib/synchronize-service/ in firmware repo

Directory structure

lib/synchronize-service/

bin/synchronized-hostapd-handler     # WiFi AP event handler

synchronize_service/

__main__.py                        # Entry point

constants.py                       # Port, paths, env config

client/http.py                     # MicroPython HTTP client

http/

server.py                        # Routes & app factory

middleware.py                    # Error handling, ETag, validation

controllers/devices.py          # REST API endpoints

controllers/ui.py               # Web UI endpoints

schemas/                        # Marshmallow schemas

views/                          # Jinja2 templates

models/

database.py                     # SQLite, type converters

migrations.py                   # Schema DDL

device.py, event.py, batch.py   # Data models

settings.py, file.py            # Settings, file upload

access_point.py                 # WiFi connectivity

6.2 Running the Service

pip install -e lib/synchronize-service/

python -m synchronize_service.__main__    # or: synchronized

On startup: creates data directory and blob storage, initializes SQLite, queries hostapd_cli list_sta for initial WiFi state, starts aiohttp server on port 9090.

6.3 API Endpoints

Method

Path

Description

GET

/devices

List all devices

GET

/devices/{uid}

Get device state, settings, pending commands

POST

/devices/{uid}

Create or update device state

GET

/devices/{uid}/events

List device events (paginated)

POST

/devices/{uid}/batch/{batch_id}/events

Batch upload events (JSON or JSONL)

POST

/devices/{uid}/files

Upload file (streamed, SHA-256 addressed)

GET

/devices/{uid}/files

List uploaded files

GET

/devices/{uid}/files/{id}/blob

Download file blob

POST

/devices/{uid}/settings

Update time table and commands

Web UI: /ui/devices (device list), /ui/devices/{uid} (device detail with settings, events, files), /ui/devices/machine (set system clock).

6.4 Database Schema

SQLite with STRICT mode. Foreign keys enforced. Auto-triggers create settings and access_point rows on device insert.

Table

Purpose

Key columns

devices

Device identity and latest state

uid (BLOB, unique), firmware_version, pump revolutions, file system stats, IP/MAC, deleted_at (soft delete)

batches

Groups event uploads

device_id (FK)

events

Device log entries

uid, type (free-text, CHECK length>0; the EventType enum has 9 values: log/state/command/settings/sample/initialization/recovery/warning/error), message, arguments (JSON), emitted_at, created_at

settings

Per-device config

device_id (unique FK). Auto-created.

time_table_entries

Sampling schedule

minutes (offset from session start), samples (count)

commands

Pending device commands

uid (4-byte random), name, arguments (JSON). Removed when acknowledged.

access_points

WiFi connectivity

connected (0/1). Auto-created.

files

Uploaded file metadata

name, size, content_hash (SHA-256)

Command names (13, the CommandName enum): prepare_session, anticoagulant_bag_punctured, anticoagulant_primed, catheter_inserted, blood_primed, start_session, reset_session, restart_device, upload_file, ping, clean_event_store (0.1.102/3, SWR-263: uid-continuous store wipe, idempotent per command uid), telemetry_on, telemetry_off (0.1.105, SWR-267: /telemetry flag + live task toggle, default OFF). (Corrected 2026-06-19 — the previously listed maintain_anticoagulant_pump does not exist anywhere in the code.)

6.5 Device Communication Protocol

Devices sync over WiFi using synchronize_service/client/http.py. Three-step flow:

Step 1 — Register / update:

POST /devices/{uid}

{ "latest_event_uid": 42, "blood_pump_revolutions": 1000,

"firmware_version": "0.1.13", "mac_address": "aabbccddeeff",

"event_batch": true, ... }

Response includes settings (time table + pending commands) and optional batch ID.

Step 2 — Upload events:

POST /devices/{uid}/batch/{batch_id}/events

Content-Type: application/jsonl

{"uid":43,"type":"log","message":"Starting.","arguments":{},"emitted_at":1234567891}

{"uid":44,"type":"command","message":"Ack.","arguments":{"command_uid":"12345678"},"emitted_at":1234567893}

JSONL (streamed, one per line) or JSON array. Command acknowledgement removes pending command from settings.

Step 3 — Upload files:

POST /devices/{uid}/files

Filename: error.txt

<binary body>

Stored content-addressed by SHA-256 hash.

ETag caching: Client sends Exclude-Match: P/"settings.time_table:{hash}". Server omits unchanged settings, responds with Exclude-Applied header.

6.6 WiFi Connectivity Tracking

Raspberry Pi runs as WiFi AP via hostapd. Script bin/synchronized-hostapd-handler registered with hostapd_cli -a receives AP-STA-CONNECTED / AP-STA-DISCONNECTED events and updates access_points table by MAC address. On startup: hostapd_cli list_sta initializes state.

6.7 Web UI

Accessible at http://fluispotter.internal/synchronize/ui/devices. Jinja2 templates.

Device list — columns are Synchronized (last sync), UID, and Status (online/offline), plus a per-row actions column. (Firmware version appears on the device-detail page, not the list.) A system-clock (“Machine”) dialog is available for field use without NTP.

Device detail — Full state. Time table editor (CSV import/export). Command queue (context-aware state commands + global: restart, ping, upload_file, reset_session). File browser with download/preview. Paginated event log (200/page).

6.8 MicroPython Client

Async Device class in client/http.py. Runs on ESP32.

device = Device(base_url="http://hub.local/synchronize", uid=uid, request=request_fn)

await device.create_or_update(latest_event_uid=42, event_batch=True, ...)

await device.batch_create_events(events)

await device.upload_file_blob(blob, name="error.txt")

await device.read()

Maintains ETag for settings caching. Data structures: Batch(id, size), Event(uid, type, message, arguments, emitted_at), Settings(time_table, commands), TimeTableEntry(minutes, samples), Command(uid, name, arguments).

Validation: firmware_version matches ^\d+\.\d+\.\d+$. mac_address = 6 bytes. ip_address = 4 or 16 bytes.

7. Software Architecture — Component Overview

7.1 System Diagram

macOS (development)                                                 ESP32-S3 (device)

+--------------------+                                             +----------------------+

| 1. Monitor App     |----------------USB serial------------------>| 4. Main Firmware     |

|    (PyQt6)         |                                             |    (MicroPython)     |

+--------------------+                                             +----------+-----------+

| 2. Control App     |----------------USB serial------------------>|          |           |

|    (PyQt6)         |                                             +----------+-----------+

+--------------------+           Raspberry Pi (field hub)          | 5. ULP Co-processor  |

+----------------------+          |    (C + RISC-V asm)  |

| 3. Synchronize       |<--WiFi--+----------------------+

|    Service            |  HTTP

|    (aiohttp + SQLite) |

|    Web UI on :9090    |

|    hostapd AP mgmt    |

+----------------------+

Five distinct software components span three hardware targets. Components 1–2 are development/test tools; component 3 is the end-user field interface; components 4–5 run on the wearable device.

7.2 Component 1 — Monitor App (macOS)

Repository path

fluispotter-firmware/monitor/

Framework

PyQt6 + matplotlib

Dependencies

pyserial ≥ 3.5, PyQt6 ≥ 6.4, matplotlib ≥ 3.7, numpy ≥ 1.24

Purpose

Read-only live telemetry dashboard for development and testing

Connection

USB serial via MicroPython raw-REPL protocol (115200 baud)

Connects to the ESP32-S3 over USB, interrupts running code, injects a monitoring script that emits JSON at 100 ms intervals. A SerialBridge daemon thread reads JSON into a thread-safe queue; a 100 ms QTimer drains the queue and updates the UI.

Tabs: Live (state machines + variable dashboard + sparklines), Tape & Valve (valve position, tape reel, reed switch), Sensors (analog/digital sensors with trend arrows and mini-sparklines), Sampling (hall event log), Log (raw JSON stream, color-coded).

7.2.1 Passive Telemetry Mode (Development Only)

The default Monitor connection mode uses raw-REPL injection, which interrupts the running firmware. For development and debugging while the firmware state machine is running, a non-interrupting passive telemetry mode is available.

How it works: A background asyncio task inside the main firmware periodically prints JSON telemetry frames to stdout (USB serial). The Monitor app opens the serial port in listen-only mode and parses the incoming JSON without injecting any code.

Enable

Create an empty file /telemetry on the device filesystem, then reset:

      rshell -p /dev/tty.usbmodem* touch /pyboard/telemetry

Disable

Remove the flag file and reset:

      rshell -p /dev/tty.usbmodem* rm /pyboard/telemetry

Flag file

/telemetry — presence checked at boot; no firmware redeployment needed

Output format

Same JSON schema as the injected board_script.py (ULP variables, GPIO states, ADC sensors, I2C sensors)

Cadence

Fast loop: 100 ms (GPIO, ADC, ULP mem32). Slow I2C sensors (VEML6035, SHT4X): 1 s.

Important: This feature is intended for development units only. Do not enable on units deployed at customer sites — the continuous serial output is unnecessary overhead and may interfere with rshell/REPL access.

7.3 Component 2 — Control App (macOS)

Repository path

fluispotter-firmware/control/

Framework

PyQt6 + matplotlib

Additional dependency

psycopg[binary,pool] ≥ 3.1 (PostgreSQL board registry)

Purpose

Active device control, test execution, and calibration

Extends the Monitor architecture with writable serial commands and injectable test scripts. The SerialBridge adds run_command() and run_command_with_output() for command injection.

Additional tabs: Pumps (AC+BP control, Power/SPR modes, time-series charts), Resistance Map (polar chart — 32 segments/rev, viridis colourmap), Encoder (tick fidelity vs 5184 reference), DAC Sweep (DAC vs SPR scatter), PID Tuner (three modes — Operational tunes the LIVE Simple-P production regulation via the per-pump *_p_step_per_pct + shared op_p_* knobs and drives the real pumps; Test bench-tunes the legacy positional PID; Simple-P (16-tick) exercises the test2 sandbox), Sensors (per-pair LED enables + live sparklines with adjustable cutoff tuning — tinted classification bands over the retrospective samples for Blood/Water/Air, AC Water/Air, Spot and Line, defaults from board_config), Parking (pump calibration), Valve-Tape (GPIO write controls), Event Log (on-board AppendOnlyStore), Board Registry (PostgreSQL).

Injectable test scripts: resistance_map_script, encoder_fidelity_script, fixed_power_run_script, dac_sweep_spr_script, pid_tuner_script (legacy PID), pid_regulate_tune_script, pid_p_regulate_script (test2 Simple-P), pid_operational_p_regulate_script (live operational Simple-P), sensors_monitor_script, park_pump_script, determine_parking_script, event_log_script.

Single-channel telemetry. The board runs one injected script at a time over the single USB-serial link; the Control App swaps scripts on tab change (live MONITOR_SCRIPT by default, a tab-specific script while a test/Sensors tab is active). Consequently only the active tab's data stream updates — passive tabs freeze until re-selected, and hardware-driving action tabs (PID Tuner, Resistance Map, DAC Sweep, Parking) are necessarily exclusive. A future option is to unify the passive monitoring tabs onto the one MONITOR_SCRIPT stream so they update concurrently.

7.4 Component 3 — Synchronize Service (Raspberry Pi)

Repository

synchronize-service/ (separate repo)

Framework

aiohttp 3.8.6, aiosqlite 0.19, Jinja2 3.1.2, marshmallow 3.20.1

Database

SQLite (async via aiosqlite)

Port

9090 (configurable via HTTP_PORT env var)

Entry point

synchronized console script (synchronize_service.__main__:main)

Storage

/data/synchronize-service (database, blobs, temp files)

End-user hub running on the Raspberry Pi. Manages Fluispotter devices in the field via WiFi AP (hostapd). Provides a JSON API for ESP32 clients and a browser-based HTML UI with real-time WebSocket updates (morphdom DOM-diffing).

Device flow: Device connects to RPI WiFi AP → POSTs registration (/synchronize/devices/{uid}) → receives settings (time table + commands) → uploads events in batches → uploads data files (content-addressed by SHA256).

Web UI pages: Device List (all devices, connection status, acquire/release/reject actions, machine time setting). Device Detail (info, 4 time-series graphs — AC SPR, BP SPR, tape pulses, fluid sensor — configurable time window, settings editor with time-table CSV import/export, command queue, paginated event log, file browser with version history).

Commands: PREPARE_SESSION, MAINTAIN_ANTICOAGULANT_PUMP, ANTICOAGULANT_BAG_PUNCTURED, ANTICOAGULANT_PRIMED, CATHETER_INSERTED, BLOOD_PRIMED, START_SESSION, RESET_SESSION, RESTART_DEVICE, UPLOAD_FILE, PING, PURGE_CATHETER.

WiFi tracking: synchronized-hostapd-handler script called by hostapd_cli -a on STA connect/disconnect events; updates access_points table via direct SQLite query.

See §6 for full API endpoints, database schema, and protocol details.

7.5 Component 4 — ESP32-S3 Main Firmware (MicroPython)

Repository path

fluispotter-firmware/fluispotter/

Runtime

MicroPython on ESP32-S3-WROOM-1-N8R8

Entry point

main.py (USB host detect → REPL or fluispotter.main())

Boot sequence: Checks for USB host via SOF interrupt (5 ms). If USB host present (development), drops to REPL for Monitor/Control app. If no host (field), runs the main firmware state machine.

Key modules:

Module

Purpose

fluispotter.py

Main state machine, WiFi, NTP sync, hub synchronisation (1 s poll — a full status POST + event upload every 5th poll ≈ every 5 s, a lightweight command GET on the others), ULP pump-control supervision (commands the ULP; the ULP runs the regulation), sensor sampling, valve/tape control, event logging

pump_regulator.py

Loads ULP binary into RTC slow memory, manages main-CPU / ULP interface

variables.py

ULP shared-memory address map in RTC slow memory. Generated per build from the ULP ELF symbols — the exact addresses shift every build, so consumers import the NAMES, never a fixed base (the old “0x50001C94+” literal was stale/incorrect).

event_store.py

Persistent event logging via AppendOnlyList on flash (8192 events)

aiorequests.py

Async HTTP client for synchronise-service communication

board_config.py

GPIO pins, I²C addresses, ADC channels, hardware constants

winbond.py

SPI driver for Winbond W25Q64 external NOR flash

Peripheral drivers (in lib/): MicroPython-DAC7574, MicroPython-MCP79410, MicroPython-VEML6035, MicroPython_SHT4X, simple-pid.

Synchronisation client (lib/synchronize-service/synchronize_service/client/http.py): MicroPython HTTP client implementing the device side of the protocol — create_or_update(), batch_create_events(), upload_file_blob(), with ETag support.

State machine: the production controller (MainController) runs an explicit state machine — INITIAL → AWAIT_CALIBRATION → CALIBRATION → AWAIT_ANTICOAGULANT_BAG → ANTICOAGULANT_BAG → AWAIT_CATHETER → CATHETER → AWAIT_SESSION → SESSION → FINISH — with the current state mirrored to ULP shared memory (main_cpu_state) and persisted to state.dat so a reset mid-SESSION resumes the schedule (SWR-230). SESSION runs three sub-states internally (IDLE / TRIGGER / DOSING). Background work runs as cooperative asyncio tasks created in run() (telemetry, reed-ON watch, AC + blood-pump cycles, fluid sampling) and torn down in finally; run_safe() wraps the whole machine and on any unhandled error appends to fatal-error.txt + the diagnostics CSV, then resets.

Watchdog model (SWR-215, rewritten 2026-07-10; supersedes the earlier “no watchdog” stance): machine.WDT(30 s) is armed once a session exists (fed by a 1 s asyncio task) so an event-loop stall becomes a reset that RESUMES the session (SWR-257/259); pre-session states run under a lenient 120 s run()-wide floor (_WDT_PRESESSION_TIMEOUT_MS) because the suspend+DTR CDC freeze is invisible to the serial sink’s latch. Bench REPL work MUST call park_wdt() first — interrupting the firmware kills the feeder and the pending reset otherwise drops the serial port mid-transfer (the recurring “Device not configured” deploy failure, root-caused 2026-07-23). The old “WDT reset is a false recovery for a latched DAC7574” concern is superseded by the crash-boot recovery in pump_regulator.start() (I2C bus recover + DAC safe-off on every non-PWRON boot). Non-blocking serial (SWR-235, rewritten 2026-07-10): ALL console writes route through serial_sink (measure-and-latch): each write is timed and the FIRST slow one (≥200 ms) trips a one-way latch — no further code path touches the CDC. Known hole (issue #53): in the host-suspended+DTR-latched regime the C-layer write NEVER RETURNS (mp_usbd_cdc_tx_strn not-connected loop has no timeout), which the sink cannot measure; the watchdog floor is the backstop, and the §7.5a boot-loop breaker below closes the loop-forever case.

ULP in-place self-recovery (fw 0.1.111, SWR-275): when the health gate finds the #36 stack canary tripped or the ULP core stuck, the firmware attempts ONE in-place recovery — unconditional I2C bus clear + DAC CH0/CH1 safe-off, BP direction re-set, fresh ULP binary load with a FULL .bss zero (clears the canary latch), SWR-257 liveness proof — then re-verifies; a power-cycle is demanded only if the recovery itself fails. Customer battery devices cannot be power-cycled. Live-proven on the bench (artificial canary trip → recovered, ULP intact). The loader also pattern-fills the top 640 B of the ULP stack gap; stack_low_water() reads the true deepest excursion (idle: 228 B; measured static worst chain: 480 B of 696 B headroom).

Actuator power-on safety (SWR-276, fw 0.1.112–0.1.114 + IMG-4): the DAC7574 powers up at zero-scale = FULL DRIVE in the inverted convention (power = 4095 − register), and the AC driver enable is biased ON board-side while GPIO9/17 both float at reset (S3 datasheet Table 2-1). Defense in depth, all commanding enables LOW + DAC HALT: (1) second-stage bootloader hook (bootloader_before_init, IMG-4) drives GPIO 9/17/8/41 LOW with pull-downs ~tens of ms after power-on; (2) boot.py at ~0.5 s (field-measured); (3) main.py repeats before the USB gate. Residual exposure = the mask-ROM window (~10–50 ms) — closable only by the hardware pull-down on the AC enable, queued for the next HW revision.

Clock-first boot ordering (SWR-295/296, issues #52/#53 — 2026-07-23): two live incidents established the clock discipline now in the registry. (1) A mid-session WDT recovery can resume BEFORE NTP completes; with the RTC at the year-2000 power-on epoch every scheduled sample reads +26 years away and the session silently flush-naps forever while event sync works (issue #52; the operator-visible symptom is “stalled updates”). SWR-295: a pre-build-epoch clock forces an immediate hub NTP sync with bounded awake-retry in EVERY state; recovery must not schedule against it and must error loudly at the retry bound. (2) The boot flow entered AWAIT_CALIBRATION before NTP, so operator-facing events carried 2000-01-01 stamps. SWR-296: WiFi first, then NTP inside the SWR-256 plausibility window, and only then the AWAIT_CALIBRATION transition / acquired report / command processing — only the Boot/INITIAL forensic chain may ever be pre-epoch. Firmware implementation of both is queued (registry DOC 1.90).

7.5a Boot-loop breaker, quiet-console latch & boot breadcrumbs (issue #53, fw 0.1.132–0.1.134): a WDT reset in a pre-session state has nothing to resume: if the underlying freeze recurs (battery + detached USB in the radio-hot acquire loop reproduced it every 1.5–4 min), the device loops forever. bootstat.py (top-level module) keeps a 6-byte protocol in RTC user memory (magic 0xB5; survives WDT/soft reset, clears on power loss): consecutive-WDT-boot streak, the previous boot’s final breadcrumb, and the live breadcrumb. Every Boot event carries crumb + wdt_streak. At streak ≥3 the breaker creates /console_quiet and serial_sink is born latched from then on (no code path touches the CDC; delete the file to restore prints; /serial_sink_off dev bypass still wins). Breadcrumb ids bracket every suspect call: 10 poll-loop top; 11/12 console write (a WDT boot reporting 11 died INSIDE a CDC write); 13/14 fast settings GET; 15/16 full event sync; 17/18 NTP; 19/20 scan; 21/22 connect; 23/24 command handling. Fork facts established en route: machine.reset() reads back as HARD (code 2), and RTC user memory is wiped by a full firmware boot (suspected overlap with the ULP’s reserved slow-mem at reload) — tolerated because the per-poll crumb rewrites re-establish the blob after every ULP load.

7.6 Component 5 — ULP Co-processor Firmware (C + RISC-V Assembly)

Repository path

fluispotter-firmware/esp32/pump_regulator/

Toolchain

ESP-IDF with ULP-RISC-V

Output

pump_regulator.bin (loaded at runtime by main firmware)

Real-time pump control running independently of the main CPU. Operates during Light Sleep. See §5.5 for detailed behaviour.

Three images from one source: main.c compiles to three binaries selected by build defines — pump_regulator.bin (operational), pump_test.bin (ULP_TEST_BUILD — full diagnostics) and pump_test2.bin (ULP_TEST_BUILD+ULP_TEST2_BUILD — the trimmed PID-Tuner image), each with its own version + address map (variables.py / test_variables.py / test2_variables.py). The ULP RTC reservation is 8176 bytes (CONFIG_ULP_COPROC_RESERVE_MEM): the operational image is boot-loaded into it, but a test image swapped in at runtime must stay under a 4592-byte ceiling (TEST2_MAX_BYTES) — above it the firmware's mem32 binary write overruns into live MicroPython RTC memory and crashes (which is why the dose/park, too large to swap, is debugged against the operational binary).

Source files:

File

Purpose

main/ulp/main.c

ULP main loop: reads Hall GPIOs, manages AC/BP state machines, writes DAC, updates shared memory, wakes main CPU on hall tick

main/ulp/dac7574.c/.h

DAC7574 driver over the ESP-IDF ULP-RISC-V RTC-I2C hardware peripheral (ulp_riscv_i2c_master_*; device at 0x4C) — the ULP does have hardware I²C; it is NOT bit-banged.

main/ulp/pulse_counter.S

Legacy stub (TODOs only, does no pulse counting) and NOT in the build — CMake compiles only main.c + dac7574.c. Encoder edges are counted in C (sample_encoder in main.c).

main/ulp/config.h

Blood-pump encoder RTCIO register macros only; included by the (unused) pulse_counter.S, not by main.c. The real GPIO map, the 0x4C DAC address and timing constants live in main.c.

main/variables.c

MicroPython dynamic-runtime stub (mpy_init) — contains no shared-var definitions. The shared vars are C globals in main.c; their addresses are emitted into fluispotter/variables.py by the build.

ULP versioning: The ULP binary carries its own version in a shared-memory constant ulp_version, packed as (major<<16 | minor<<8 | patch). The main CPU reads it via mem32 and includes it in telemetry as ulp_ver. The ULP version is independent of the MicroPython firmware version and is bumped on every ULP C change.

Automated deployment: The deploy_ulp.py script in the repository root automates the full ULP build-deploy-verify cycle. It runs 14 sequential steps: (1) ninja sub-build, (2) objcopy ELF→BIN, (3) address regeneration from ELF symbols, (4) verifies consumer files import addresses by NAME — scans Monitor/Control for hardcoded 0x5000xxxx literals and aborts if any are found (the old “atomic hex-address remap” was deliberately removed; addresses live only in variables.py), (5) firmware version stamp, (6) binary deploy via mpremote, (7) Python file deploy, (8) stale .mpy detection & removal, (9) state file wipe, (10) board reset with USB re-enumeration wait, (11) RTC memory verification of the entire immutable .text+.rodata region against what was loaded (the old 16-byte reset-vector check was insufficient — 16 bytes pass even with a stale binary), (12) ULP + firmware version readback, (13) telemetry sanity check (INITIALIZED, DAC values, state words, pump-running flags), (14) binary-swap self-test (loads + runs a function in each binary). The script aborts on any failure.

Shared memory interface: All inter-processor communication via RTC slow memory. Main CPU writes targets (DAC values, running flags, PID gains); ULP writes status (state words, hall switches, pulse counts, elapsed times). Monitor/Control apps read these addresses via machine.mem32 in injected board scripts.

7.7 Repository Structure

fluispotter-firmware/                    Main firmware repo

main.py                               Boot entry (USB detect / REPL vs firmware)

deploy_ulp.py                          Automated ULP build/deploy/verify (13 steps)

fluispotter/                           [4] ESP32 main firmware (MicroPython)

esp32/pump_regulator/                  [5] ULP co-processor (C + RISC-V asm)

monitor/                               [1] Monitor app (macOS, PyQt6)

control/                               [2] Control app (macOS, PyQt6)

lib/                                   MicroPython libraries + sync client

launchers/                             macOS .app bundles

docs/                                  This specification document

synchronize-service/                     [3] RPI hub service (separate repo)

synchronize_service/                   aiohttp app, models, views, assets

bin/synchronized-hostapd-handler       hostapd event hook

7.8 Flashing the Firmware Image (MicroPython)

The device firmware ships in three layers, each deployed by a different mechanism. Reflashing the MicroPython image is the heavyweight operation and is rarely needed — a normal application fix is a serial .py deploy, and a pump-controller change is a deploy_ulp.py run. Reflash the image only to change the MicroPython runtime itself or its sdkconfig (for example the watchdog / brownout settings baked into fw 0.1.98 below).

⚠ CRITICAL (issue #54, discovered 2026-07-23): the IMG-4 image (v1.25.0-12, 2026-07-16) freezes its own main.py, and the frozen copy runs INSTEAD of the filesystem /main.py — proven live: the firmware boots identically with /main.py renamed away. Every /main.py deploy since IMG-4 has been silently dead code (the deploy tooling checksum-verifies the file and reports success — a false-confidence loop). Until an image rebuild un-freezes it (IMG-5, one-change + boot-prove discipline), ALL early-boot code changes go in fluispotter/fluispotter.py module level (the earliest file proven to execute every boot; the fw 0.1.134 bootstat hook uses an idempotent guard so a future de-frozen image is safe). The “FS copies shadow the frozen ones” rule in the layer table holds for package modules — NOT for main.py itself on this fork.

Layer

What it is

How it is deployed

MicroPython image

Bootloader + partition table + MicroPython runtime + the frozen core of the app (per manifest.py). One .bin, built in CI.

esptool write-flash 0x0 (this section).

Application .py

The loose fluispotter/*.py on the littlefs filesystem. FS copies shadow the frozen ones, so an ordinary code fix never needs a reflash.

Serial raw-REPL (output/deploy_firmware_serial.py).

ULP binary

The RISC-V pump-controller image + its variables.py address map.

deploy_ulp.py (see §7.6).

Building the image (CI). Push a v* git tag → .github/workflows/build.yml builds on a GitHub ubuntu-latest runner (ESP-IDF + the fluispotter/micropython fork, BOARD=ESP32_GENERIC_S3, BOARD_VARIANT=SPIRAM_OCT, FROZEN_MANIFEST=manifest.py) and publishes fluispotter-<tag>.bin as a GitHub Release asset. The private submodules (the fork, ESP-IDF, and the vendored MicroPython libraries) are cloned via the SUBMODULE_TOKEN repository secret — a classic PAT with repo read scope, mapped onto the .gitmodules SSH URLs with a global url.….insteadOf rewrite. The historical fluispotter-bot SSH deploy key is no longer valid on GitHub, so a live PAT is required for every build.

Flashing. Download the release fluispotter-<tag>.bin, then, with the board in ROM download mode (below):

python3 -m esptool --chip esp32s3 --port <PORT> \

--before no-reset --after hard-reset write-flash 0x0 fluispotter-<tag>.bin

The image is written at offset 0x0 (bootloader + partition table + app). This overwrites only the app region; the littlefs vfs partition at 0x200000 (6 MB) — which holds the loose fluispotter/*.py, the ULP binary, board_config/board_overrides.json and state.dat/time-table.dat — is preserved across the swap. esptool hash-verifies the write.

Entering ROM download mode (the S3 procedure — verified 2026-07-09). esptool cannot put the ESP32-S3 into download mode by itself over the native USB — its DTR/RTS reset gives Failed to connect: No serial data received, and repeated attempts wedge the running board. Use one of these instead:

•	
      Physical Boot1 button (the robust method — works even when the flashed image is broken and the REPL is dead). Boot1 is the GPIO0 boot-strapping button (§4.7 BOOT_MODE); this board has no reset/EN button. Sequence: power OFF → press and hold Boot1 → power ON while holding → keep holding ~5 s → release. The mask-ROM (independent of the flashed firmware) enters download mode and enumerates as the USB-Serial-JTAG device VID 0x303A / PID 0x1001, and the port stays up steady. (A port that flaps in and out instead means the app is boot-looping, not stable download mode — hold Boot1 through the whole power-on.)

•	
      machine.bootloader() from a live REPL — only if the app already boots and the REPL is reachable; reboots straight into ROM download mode.

This board has no auto-reset: esptool --after hard-reset pulses the RTS pin, which is not wired to EN here, so it does not restart the board. After a write-flash the board stays parked in download mode — you must hand power-cycle it (Boot1 released) to actually boot the freshly-flashed image. Read the USB PID to know the state: 0x1001 = ROM/download mode (esptool can talk); 0x303A:0x4001 = MicroPython native USB-CDC = the app is running (esptool cannot talk).

Full-image backup / restore. No flashable full image is kept in any release (the restore-fw0.1.78 release holds only the ULP .bin files). Before flashing, dump the whole 8 MB via the same download-mode entry (--before no-reset because the board is already in download mode):

python3 -m esptool --chip esp32s3 --port <PORT> \

--before no-reset --after hard-reset read-flash 0x0 0x800000 backup.bin

Verify a good dump: first byte 0xE9 (ESP32 image magic); littlefs / MicroPython / fluispotter signatures present. Restore is an exact rollback (the dump includes the littlefs FS = the .py/ULP/config/state): write-flash 0x0 backup.bin.

A flashed image that will not boot — NOT a power or hardware problem. The failure signature — no INITIAL indicator-LED flash + the spot/line sensor LEDs stuck ON + no USB enumeration at all + no hub sync — means the flashed image itself does not boot (a bad/unvalidated runtime or sdkconfig). It is not inadequate supply and it is not a dead board (standing project rule: it is never the hardware). Do not diagnose this as a brownout or a weak/current-limited supply. Recovery is deterministic: force ROM download mode with Boot1 (above) and restore a known-good full-flash image, then hand power-cycle — the board comes straight back (proven 2026-07-09, when a rebuilt image would not boot and was rolled back). The specific reason a given build fails to boot is determined off the board from its image + build provenance (offline forensics on the .bin), never by swapping supplies.

Operational note — the PPK2 is not the board. When a Nordic Power Profiler (PPK2) is inline, it enumerates as a separate serial device: Nordic Semiconductor VID 0x1915, versus the board’s Espressif VID 0x303A. output/_serial_io.find_port VID-filters the PPK2 out, but esptool auto-detect does not — always pass an explicit --port. Never restart the PPK2 power-logger into a live DUT: restarting it resets the PPK2 and cuts the device’s power (a real interruption, unrelated to the image not booting).

7.9 Building a New Firmware Image — the Verified Method

This is the exact procedure executed and boot-proven on 2026-07-09 (images v0.1.99-base and v0.1.99, the interrupt-watchdog fix for issue #44). Every step below was performed at least once; nothing here is theoretical. It exists because its violation — delivering a config tweak by bumping the runtime submodule to the fork tip, which silently swapped the entire MicroPython runtime a dev-cycle forward — produced the non-booting v0.1.98.

Rules (standing): (1) one sdkconfig change per image, never stacked; (2) the runtime base is always an explicit, proven commit pin — never a branch tip; (3) every image passes the pre-flash gates before it touches the board and the boot-proof before it is trusted; (4) a restorable full-flash backup exists before flashing (§7.8).

Step

Action (as executed)

0. Backup

Confirm a verified full-flash backup of the current working image exists (§7.8 read-flash). This is the rollback for every step below.

1. Fork branch

In fluispotter/micropython, create a branch off the proven base commit (2026-07-09: fluispotter-v1.25-int-wdt-1000ms off 1d140090, the v1.25.0-esp32 tip the firmware repo already pinned) and add one commit containing the single sdkconfig.board change (the INT_WDT commit is 3c8d3909, +5/−0: four comment lines + CONFIG_ESP_INT_WDT_TIMEOUT_MS=1000). Verify the commit diff touches only that file. Never base on master (the fork's master is a v1.26-preview line — the v0.1.98 brick). The commit must be on a pushed branch so CI's git submodule update can fetch it.

2. Pin bump

In the firmware repo, move only the runtime/micropython gitlink to the new commit: git update-index --cacheinfo 160000,<sha>,runtime/micropython; commit; verify with git ls-tree HEAD runtime/ that runtime/esp-idf is unchanged (57fa26be, the v5.2.2-ulp pin that matches what MicroPython v1.25.0 requires).

3. CI build

Tag v* and push the branch + tag → build.yml builds (ESP32_GENERIC_S3 / SPIRAM_OCT, ~9 min) and publishes fluispotter-<tag>.bin as a Release asset.

4. Pre-flash gates

On the downloaded .bin, BEFORE it touches the board: Gate 1 — strings -a <bin> | grep 'MicroPython v1.' MUST report exactly the intended pin (e.g. v1.25.0-8.g3c8d3909), not an unintended runtime (this single gate would have stopped v0.1.98, which read v1.26.0-preview.151). Gate 2 — esptool image-info: header identical to the proven image (Flash DIO/80m/8MB, entry point, chip-rev) and checksum + validation hash valid.

5. Flash

Enter ROM download mode with Boot1 (§7.8; PID 0x1001, esptool talks), then write-flash 0x0 <bin> (littlefs at 0x200000 preserved; esptool hash-verifies). The board stays in download mode afterwards — the RTS “hard reset” does nothing on this board.

6. Boot-proof

Plain power-cycle, Boot1 untouched. Acceptance = ALL of: the INITIAL (blue) LED at boot + USB re-enumerates as native CDC 0x303A:0x4001 + hub sync resumes. Only a fully boot-proven image is trusted; a held Boot1 lands back in 0x1001 download mode (harmless — power-cycle again without it).

7. Sequence

If the pipeline has never produced a bootable image from the intended base, build an Image 0 first: the base pin with zero config change (2026-07-09: v0.1.99-base at 1d140090, boot-proven), then the one-change image on top (v0.1.99 at 3c8d3909, boot-proven). A failure then attributes unambiguously: Image 0 fails → the base/pipeline; Image 1 fails → the single change.

8. Rollback

Any failure at step 6: Boot1 → download mode → write-flash 0x0 <backup> → plain power-cycle. Proven same-day (the v0.1.98 recovery).

Executed instances (both boot-proven 2026-07-09): v0.1.99-base — runtime v1.25.0-7.g1d140090, no config change, the pipeline-proof image. v0.1.99 — runtime v1.25.0-8.g3c8d3909, + CONFIG_ESP_INT_WDT_TIMEOUT_MS=1000 only: the delivered fix for the intermittent post-dose interrupt-watchdog reboots (issue #44). Brownout remains deferred as its own future image (it would only pin the existing esp-idf level-7 default).

Runtime sdkconfig changes are image-level and must be built + boot-proven ONE at a time. Changing the ESP-IDF sdkconfig (interrupt-watchdog, brownout, PSRAM, console, …) requires rebuilding the whole MicroPython image — it cannot be a loose .py deploy. The two intended safety changes are CONFIG_ESP_INT_WDT_TIMEOUT_MS=1000 (raises the interrupt-watchdog margin from the 300 ms default to 1000 ms — the fix for the intermittent post-dose “WDT code 3” reboots, the interrupt watchdog firing on a transient >300 ms interrupts-off window in the Wi-Fi / light-sleep path, GitHub issue #44) and enabling the brownout detector (CONFIG_ESP_BROWNOUT_DET).

Incident — v0.1.98 (2026-07-09) did NOT boot and was rolled back. A single fork rebuild bundled both sdkconfig changes on top of a year-newer MicroPython + esp-idf runtime, and the resulting image would not boot (no USB enumeration at all). It was recovered by restoring the known-good full-flash backup. Lessons now standing: (1) change one sdkconfig setting per image (interrupt-watchdog first; brownout — least critical — last), never stacked, so a boot failure is attributable; (2) every separately-deliverable unit (image/runtime, app, ULP) carries its own version + build provenance stamp; (3) a freshly built image is boot-proven on the bench (INITIAL LED + USB 0x4001 + hub sync) before it is trusted or endurance-tested. The interrupt-watchdog change was re-delivered under this discipline and boot-proven the same day (v0.1.99, step table above); brownout remains deferred.

7.10 Versioning & Image Provenance (task #114, SWR-268)

Every separately-deliverable unit is versioned (owner directive 2026-07-09). The units and their version homes:

Unit

Version home

Discipline

MCU firmware (.py)

fluispotter/version.py (MAJOR.MINOR.PATCH + changelog line + build date)

bump on EVERY change; reported to the hub every sync

ULP binary

ULP_VERSION in main.c; reported via RTC shared mem → hub device page

rebuild + regen variables.py + deploy checklist

MCU image (runtime + sdkconfig)

verification/images.py IMG registry — per image: fork tag, runtime commit hash, base image, the ONE config delta, build + boot-proof dates, status (BUILDING→BUILT→BOOT_PROVEN→ACTIVE→SUPERSEDED/BRICKED)

one sdkconfig/C change per image (§7.9); exactly one ACTIVE image; a CI/verify gate REFUSES a runtime/micropython gitlink that matches no registered buildable/active image (the silent submodule bump that shipped the v0.1.98 brick is structurally impossible)

Image fingerprint at runtime

the boot config event carries os.uname().version (embeds the fork commit, e.g. v1.25.0-8.g3c8d3909)

every boot in every hub log is traceable to an IMG registry row

Monitor / Control apps

app_version.py each; shown in window titles

bump + changelog on every change

Hub service

synchronize_service/version.py; shown in every page footer

bump + changelog on every change

Requirement registries

DOC_VERSION + CHANGE_HISTORY rows in requirements.py / hub_requirements.py / qc_requirements.py

locked-requirement immutability (SWR-511) + change history on every edit (SWR-512)

Registered images (as of 2.24): IMG-0 (pre-registry b10038af base, SUPERSEDED), IMG-1 (v0.1.99-base 1d140090, pipeline proof, SUPERSEDED), IMG-2 (v0.1.99 3c8d3909, INT_WDT=1000 ms, ACTIVE — field-proven 20/20 + 42/42 + 20/20), IMG-3 (v0.1.107 e9d9c06, CDC zero-progress break, BUILDING).

IMG-4 (tag v0.1.114, runtime a832022, ACTIVE since 2026-07-16): single change on IMG-3 — the SWR-276 bootloader hook (plus the tinyusb ==0.19.0~3 reproducibility pin, zero runtime delta). Boot-proven: fingerprint v1.25.0-12.ga83202261, littlefs intact. IMG-3 SUPERSEDED (its CDC fix and all prior config carry forward in the lineage IMG-1 → 2 → 3 → 4).

8. Open Items & Design Log

8.y Design log 2026-07-20→23 — #48 conviction, deque truth, clock discipline, boot-loop class, frozen main.py

#48 wild motors: solved. Global RTC-pad register-clock gate convicted by ULP-side capture; tick-top DEFEND_CLOCK() proven on the TestOp dev track (details §5.5). Deep-chain call shapes are FORBIDDEN (stack-frame growth invisible to static headroom; dev.0.4 tripped 5/5, dev.0.5 wedged sessions via the #49 unbounded AC-first wait). Operational 1.16.0 roll-in awaits the headroom budget decision (552 B vs the 768 B gate; baseline 1.15.0 already sat at 568). En route: hub 1.14.1 (dev-version schema 422 outage), fw 0.1.128 pad forensics, 0.1.129 digital-residue retire (refuted as trigger by trial).

Deque: registration complete, stamps fixed. Owner-requested per-crossing audit proved zero tick loss (120/120 + 70/70 across 12 reboots); three stamp-corruption modes convicted and closed in fw 0.1.131 (§5.6.3, issue #51 CLOSED with live acceptance). sampled_at terminology fixed: the stamp marks when the blood entered the catheter at the tip.

Clock discipline (SWR-295/296, issues #52/#45). Recovery can beat NTP and strand a session on a year-2000 clock (silent flush-napping, samples +26 years away) — almost certainly old #45’s mechanism; AWAIT_CALIBRATION entered pre-NTP pollutes the operator record with 2000-01-01 stamps. Both now registry requirements (§7.5); firmware implementation queued.

#53 battery boot loop / #50 dark boots: one freeze class, three faces. A USB-CDC write in the suspended+DTR regime never returns (no C-layer timeout) and the serial sink cannot measure it; the 120 s pre-session WDT converts the freeze into a reset; pre-session there is nothing to resume → boot loop (reproduced every 1.5–4 min on battery in the acquire state); pre-WDT-arming the same freeze would present as the multi-hour #50 dark boots; in-session the 30 s WDT self-heals it (#46’s battery fire). Countermeasures shipped (fw 0.1.132–0.1.134): boot-loop breaker + quiet-console latch + breadcrumbs (§7.5a); final conviction of the frozen line pends the instrumented battery run.

Frozen main.py (issue #54) — operational landmine defused. IMG-4 freezes main.py; the filesystem copy NEVER executes (§7.8 warning). Explains why the 0.1.132/0.1.133 breaker was inert (impossible telemetry: WDT boots with streak 0). Also established: machine.reset() → HARD(2) on this fork; RTC user memory persists across resets but is wiped by a full boot (ULP slow-mem reload overlap suspected); bench rule — park_wdt() before ANY interactive REPL work, or the 120 s floor resets the board mid-transfer (this was the recurring deploy port-killer).

Registry/process. SWR-295/296 added (DOC 1.89/1.90, owner instructions); all 28 open GitHub issues now carry a “Seen on” version footer; issues #48/#49/#50/#52/#53/#54 filed with full version/image sets; #51 closed with acceptance evidence; battery campaign step 1 (20 samples, one 2+8 flush before the last, zero-reboot bar) staged and gated on #53.

8.x Design log 2026-07-18 — battery proof, time integrity, multi-spot campaign

Battery endurance, run to power failure (2026-07-16→18). 1100 mAh: 44 h 10 m at ~24.9 mA average (22/22-sample session, 38.4 h, zero warnings/errors/reboots, avg timing delta +10.9 s; then 5.7 h post-FINISH flush). Low-battery death characterized: NO brownout-reset storm — a silent stop during a lightsleep nap, last record an orderly flush-cycle park+disable (line sealed, actuators off — the safe posture by construction); motor and RTC healthy to the end; ZERO data loss (final sync succeeded after the last log). Early-warning signature: sync TimeoutErrors confined to the final ~5 h (Wi-Fi transmit degrades first under sag). OPEN owner decision SWR-287: hardware battery telemetry vs firmware timeout-burst warning vs deliberate low-battery safe shutdown; note that an unattended finished session burns the battery at full flush duty. Full report: docs/battery_test_2026-07-18.md; evidence archived with SHA-256 manifest.

Time integrity (SWR-282..286, hub HUBR-316..319). Chain verified against code: MCP79410 RTC ← SNTP from the hub (plausibility-window gated, never stepped mid-session while sane) ← hub ntpd ← pool when uplinked; heartbeat = plausibility only. Gaps closed/planned: quantified RTC-vs-NTP offsets at session start+end (SWR-283, fw 0.1.116) with the ±30 s in-session limit and report DEVIATION flagging (SWR-284); Time 0 redefined as the operator’s start_session submit (SWR-285 — fw ≤0.1.115 measurably anchors at device receipt, +3 s live; the hub already stamps arguments.time0); schedule immutable after session start (SWR-282, locked-in owner decision, field-proven by the 22/24 incident); persistent device session number (SWR-286) — uid+number = globally unique session id.

Multi-spot campaign findings (3 spots per timepoint, 2026-07-18). Two real defects found by owner bench observation: (1) _dosing_tick step 5 restarts the pump in continuous run after EVERY dose — between set doses this wastes ~a full revolution and skips tube segments (SWR-288: final-dose-only restart; dose k+1 after exactly one ledger revolution — sets dose CONSECUTIVE segments back-to-back, ~147° repositioning). (2) bp_hall_revolutions increments ONLY in BP_RUNNING — each dose cycle’s park crossing advances the tube one segment invisibly to the deque ledger (SWR-289: count crossings in EVERY ULP state + an ULP-owned crossing-timestamp ring so the CPU deque becomes a read of ULP truth, also eliminating CPU-poll timestamp batching). Dose 1 of well-separated samples was always correct (gate windows contain only counted RUNNING revolutions — why all audits were clean). Reporting: spots and timepoints become distinct everywhere (SWR-290; settings ack gains per-entry samples counts, per-dose events gain their own timestamps).

Pre-session sample (SWR-291, planned fw 0.1.116). In AWAIT_SESSION the operator can command a sample that logs exactly like scheduled ones: scheduled_at = submit time (the SWR-285 stamp), the SAME deque gate, dose set, flush resumes, no state change; events carry sample=-1 + pre_session=true (excluded from progress counters, dedicated report row; report window widened to the session arc). Shares one commanded-sample routine with in-session sample_on_demand (SWR-280).

Report signing (hub side, for cross-reference). Session reports are dual-signed at generation: SHA-256 manifest + detached OpenSSH signature (per-hub ed25519), and an embedded certificate signature chaining to the Fluispotter Root Certificate Authority (RSA-3072, 20-year validity; encrypted root in the firmware repo, per-hub keys only on hubs). Verification material is served from the hub’s device page; one-command verifier security/verify_report.sh. Details: Hub spec §11.

8.1 Open Items

#

Item

Section

Priority

OI-01 ✓

VEML sensor — RESOLVED (see §4.1)

4.1

High

OI-02 ✓

Stroke volume — RESOLVED (see §5.1)

5.1

High

OI-03

Motor voltage variant confirmed: 12 V

2.1

High

OI-04

Encoder supply voltage and wiring

4.2

Medium

OI-05

DC-DC converter selection and efficiency

3.2

Medium

OI-06 ✓

MCU selection — RESOLVED (see §4.3)

4.3

Medium

OI-07

Runtime bench measurement

3.3

Medium

OI-08

Housing dimensions and material

5.3

Low

OI-09

Regulatory strategy (CE/FDA)

5.3

Low

OI-10

Tape-reel timeout is _TAPE_REEL_TIMEOUT_MS = 5000 (5 s) in code for both the smear and advance windows, but the spec previously documented 20 s. Confirm 5 s reliably smears + advances the tape, OR restore 20 s in the firmware. (Surfaced by the 2026-06-19 spec-vs-code audit.)

5.7

High

OI-11 ✓

RESOLVED — the §3.3 power budget is re-derived for the actual 240 MHz CPU (~165 mW / 50 mA active; runtime ~17.7 days; the motor still dominates). An optional boot-time 80 MHz downclock is noted in §3.3/§4.3.

3.3 / 4.3

Medium

OI-12

Light-Sleep wake source / wake rate (Option A, §5.6.1): the main loop naps with machine.lightsleep(60 s backstop) expecting the ULP hall/stall/fail interrupt to wake it early (as §5.6/§4.3 assume). Confirm on the bench (PPK2) that the ULP interrupt actually cuts the timed nap short, measure the effective wake rate, and reconcile idle power against the §3.3 budget — noting the blood pump runs continuously through a session, so the motor dominates session power and the CPU nap mainly helps FINISH and motor-off idle. Complete the paused PPK2 A/B (0.1.78 vs 0.1.80) idle-power comparison.

5.6.1 / 3.3

High

8.2 Session Log

Date

Topic

Key question / outcome

2026-07-20→23

#48 conviction & fix, deque audit, clock discipline, #53 boot-loop class, frozen main.py

See §8.y design log. fw 0.1.126→0.1.134; ULP op 1.15.0 + TestOp dev.0.1–0.6; hub 1.14.1; registry DOC 1.88→1.90 (SWR-294/295/296); issues #48–#54 filed, #51 closed.

2026-04-11

Software Architecture Overview

New §7: all five software components documented (Monitor, Control, Synchronize Service, ESP32 firmware, ULP firmware). System diagram, module inventories, dependency lists, and repository structure added. Informed by code analysis of both fluispotter-firmware/ and synchronize-service/ repos.

2026-04-11

ULP Hall-Wake + Lightsleep Session Loop

ULP main.c: hall edge detection in BP_RUNNING, bp_hall_revolutions counter, main_cpu_state as linker-managed variable. Session loop: machine.lightsleep() + ULP hall-tick wakes. _on_hall_tick() restores deque from ULP counter. External RTC for all time checks. Synchronize-service documented (§6).

2026-03-30

DevBoard C03191 Indexing

Schematics/BOMs indexed. MCU, VEML6035, Hall sensors confirmed. GPIO table added. OI-01, OI-06 resolved.

2026-03-26

Faulhaber motor resistance

Internal resistance queried; VEML pinout pending.

2026-03-11

Stroke Volume Calculation

Blood pump stroke volume queried.

2026-03-10

LS14500 Battery Runtime

Runtime calculated; 12 V motor + 324:1 gearbox confirmed; DC-DC ~2 V; encoder 5 % duty.

9. Source Documents & References

Document

Description

Status

EN_1512_SR_IE2-8_DFF.pdf

Faulhaber 1512 datasheet

On file

Saft LS14500 cell.pdf

Battery datasheet

On file

VEML6035 datasheet

Optical sensor register map

TBD

Onshape blood pump model

Parametric variables (§5.1.2)

Variables extracted

lib/synchronize-service/README.md

Synchronize service documentation

On file (v1.7)

esp32/pump_regulator/ULP_BUILD.md

ULP build and deployment

On file

— End of Document —