FLUISPOTTER

Software Requirements Specification

Version 1.5  |  2026-06-15

CONFIDENTIAL — Internal Use Only

Fluispotter ApS

sten@fluispotter.com

Contents

1. Document Overview

2. ULP Pump Regulation

3. Main-CPU Firmware (MicroPython)

4. Desktop Applications (Monitor & Control)

5. Synchronize Service (RPi Hub)

6. Build, Deploy & Process Gates

7. Verification & Traceability

1. Document Overview

This document is the testable software-requirements specification for the Fluispotter device software: the ULP pump regulator, the main-CPU MicroPython firmware, the Monitor and Control desktop applications, the RPi synchronize service, and the build/deploy process gates. Requirements are derived from the Device Specification Repository (v2.11), the project rule notes/checklists, and the implemented code.

Each requirement is a single testable shall-statement with a stable ID and a verification class: SW (verifiable headless — unit test, static/source analysis, parsing) or MECH (requires the physical board / pump mechanics — bench/HIL stage). This document is generated from the machine-readable registry verification/requirements.py, which the automated verification framework also uses for test tagging and coverage reporting — the document, the tests, and the report cannot drift apart.

Field

Value

Document

Fluispotter SW Requirements

Version

1.5

Date

2026-06-15

Requirements

82 total — 70 SW, 12 MECH

Source of truth

verification/requirements.py (this file is generated)

Change Log

Version

Date

Author

Change Summary

1.5

2026-06-15

Claude Code

Added SWR-224..229: a boot-loaded, firmware-independent board-config overlay (/board_overrides.json) for the 25 tunable calibration keys, with a hardware-key guard, an FNV-1a config hash, a boot event_store log + change detection (vs the persisted hash), a lock-to-freeze hook (/board_config.lock), and accessor-routed consumers. Values change without a new firmware version; applied once at boot (no runtime mutation).

1.4

2026-06-15

Claude Code

VEML6035 SPOT/LINE designations corrected — they were reversed in software. The SPOT receiver is physically at I2C 0x3a and the LINE receiver at 0x29 (board_config now defines VEML_SPOT_ADDRESS=0x3a + VEML_LINE_ADDRESS=0x29; SWR-125 + the board_config-constants requirement updated).

1.3

2026-06-15

Claude Code

BP proportional gain default raised 3 → 6 (SWR-105, SWR-123) to match operational ULP 1.3.2. Bench-validated 2026-06-15: at step=6 the blood pump converges 16 s/rev stall-free (hall-interval mean 15.96, motor_stop events 0) and tighter than 3; 7 still oscillates so remains excluded.

1.2

2026-06-15

Claude Code

SWR-122/123 (AC/BP bench convergence) now specify the TRUE SPR metric: the mean board-side time between consecutive hall revolutions, not the encoder-batch sub-sample (1/324 rev) that is dominated by within-revolution peristaltic ripple. The HIL harness computes the hall-interval SPR board-side and the bench test grades on it.

1.1

2026-06-12

Claude Code

SWR-309 reworked to the fluid-sensor STATE-BOUNDS classification model: per-state [lower, upper] bounds with the corrected signal ordering Blood < Air < Water/AC (between-bounds readings = TRANSITION) plus the four directed transition values (Air→AC, AC→Blood, Blood→Air, Air→Blood) as host-side tunables. Supersedes the ordered-cuts model seeded from board_config (whose comments encode the old, incorrect Blood < AC < Air ordering; firmware thresholds pending re-derivation from this tuning).

1.0

2026-06-12

Claude Code

Initial release: 76 testable requirements extracted from the device specification (v2.11), project rules/checklists, and the implemented firmware, ULP, and desktop-app code; each classed SW (headless) or MECH (bench). Registry is the machine-readable source of truth for the verification framework.

2. ULP Pump Regulation

The ULP-RISC-V coprocessor: Simple-P 16-tick regulation, the legacy positional PID (dosing/parking), stall and no-start interventions, parking, dosing, and the bench convergence gates.

ID

Title

Requirement

Class

Source

SWR-101

Simple-P error computation

The operational regulator SHALL compute the speed error as err% = sign(meas − target_ticks) × (|meas − target_ticks| × 100 / target_ticks), clamping meas to 4 × target_ticks before the divide (32-bit overflow guard) and flooring target_ticks at 1.

SW

spec §5.5; main/ulp/main.c run_simple_p()

SWR-102

Simple-P deadband and nudge

The regulator SHALL apply no power change when |err%| ≤ 1 and otherwise apply pwr += err% × step_per_pct with err% clamped to ±op_p_max_err_pct (default 25), continuing regulation after saturation (never halting on speed error).

SW

spec §5.5; main/ulp/main.c run_simple_p()

SWR-103

Power clamp and DAC mapping

Accumulated power SHALL be clamped to [op_p_min_pwr=64, op_p_max_pwr=4095] and written as DAC = 4095 − pwr, so the steady-state regulation DAC register never reaches DAC_HALT (4095).

SW

spec §5.5; main/ulp/main.c run_simple_p(); bug1 rule

SWR-104

Start-power seeding (bumpless)

On the first regulated update after a (re)start (first_update), the regulator SHALL seed pwr = start_power — captured at IDLE→RUNNING as DAC_HALT − preset DAC — ignoring the kick/restart DAC value.

SW

spec §5.5; main/ulp/main.c run_simple_p()/seed

SWR-105

Per-pump gain defaults

The operational ULP SHALL compile in step-per-percent defaults ac_p_step_per_pct = 7 and bp_p_step_per_pct = 6 as separate CPU-writable shared variables.

SW

spec §5.5; main/ulp/main.c globals

SWR-106

Shared Simple-P knob defaults

The ULP SHALL expose CPU-writable knobs with baked-in defaults op_p_max_err_pct=25, op_p_min_pwr=64, op_p_max_pwr=4095, op_p_pulse_count=16, op_p_skip_max=6, op_p_skip_good=4, op_p_start_escalate_n=3.

SW

spec §5.5; main/ulp/main.c globals

SWR-107

Skip-ladder level mapping

skip_for_level SHALL map level 0 to 0 skipped invocations and level N (1..SKIP_LEVEL_MAX=9) to 2^N skipped invocations (max 512).

SW

main/ulp/main.c skip_for_level()

SWR-108

Skip-ladder climb and panic rules

The %-error skip ladder SHALL climb one level after op_p_skip_good consecutive in-band samples (|err%| ≤ 1), cap at op_p_skip_max, and reset to level 0 on a panic-sized error (|err%| ≥ op_p_max_err_pct).

SW

spec §5.5; main/ulp/main.c adapt_skip_pct()

SWR-109

Hall counting through skip

Hall-revolution counting SHALL execute on every ULP tick including skipped (decimated) ticks: the AC hall read SHALL precede the skip-counter early-return in ac_tick, matching the BP path.

SW

spec §5.5 hall-through-skip; main/ulp/main.c ac_tick/bp_tick

SWR-110

Dual stall/measurement deadlines

Encoder sampling SHALL use two deadlines: a short first-edge stall window (calc_batch_timeout_us, ≥ 50 ms floor) and a longer measurement window (calc_meas_timeout_us = 4 × expected 16-edge time + 200 ms), so a dead encoder is detected fast while a slow batch can complete.

SW

spec §5.5; main/ulp/main.c calc_batch/meas_timeout_us

SWR-111

No-start kick escalation

While a started pump has produced no encoder pulse, the ULP SHALL kick at DAC_START (register 1000) per zero-pulse batch and from the op_p_start_escalate_n-th failed kick escalate to full power (DAC_MAX_POWER, register 0), setting MOTOR_START_FAIL_AC/BP once per spin-up episode.

MECH

spec §5.5 interventions; main/ulp/main.c regulate_pump()

SWR-112

Motor-stop event bit map

motor_stop_events SHALL encode bit0=AC stall (1), bit1=BP stall (2), bit2=AC start-fail (4), bit3=BP start-fail (8).

SW

main/ulp/main.c event #defines

SWR-113

Running-stall recovery sequence

On an unintended stop (pump had started, then a zero-pulse batch) the ULP SHALL halt the DAC, pause 200 ms, restart at full power, wait one restart pulse, re-seed the regulator to start_power, and set MOTOR_STOP_AC/BP — exactly one event per stall episode, captured even when the restart succeeds.

MECH

spec §5.5 interventions; main/ulp/main.c recovery branch

SWR-114

ULP version packing

Each ULP binary SHALL publish its version in the shared uint32 ulp_version packed as (MAJOR<<16)|(MINOR<<8)|PATCH, versioned independently per binary (operational / test / test2) via the ULP_TEST_BUILD/ULP_TEST2_BUILD #if ladder.

SW

main/ulp/main.c version block; deploy checklist

SWR-115

DAC convention constants

The ULP SHALL define DAC_HALT=4095 (motor stopped), DAC_START=1000 (spin-up kick ≈ power 3095), DAC_RESTART=0 and DAC_MAX_POWER=0 (full power), POWER_FLOOR=64 — the inverted convention power = 4095 − register.

SW

dac_power_convention rule; main/ulp/main.c #defines

SWR-116

Legacy PID confined to dosing/parking

run_pid (positional integer-tick PID with output-saturation anti-windup, error clamp ±16384, integral clamp [64<<16, 4095<<16]) SHALL be invoked only by the BP_DOSING and BP_PARKING state machines; steady-state regulation SHALL use run_simple_p.

SW

spec §5.5; main/ulp/main.c call graph

SWR-117

Stack-canary fail-safe

On every invocation the ULP SHALL verify the top-of-BSS canary (0xC0FFEE5A) and on mismatch latch ulp_stack_overflow=1, clear both pump-running flags, and write DAC_HALT to both motor channels before returning.

SW

main/ulp/main.c canary check

SWR-118

Test-ID dispatch partition

TEST_DETERMINE_PARKING (3) SHALL dispatch only in the operational binary; encoder/res-map/fixed-power/dac-sweep tests (1/2/4/5) only in the test binary; PID/Simple-P tuners (6/7) in both test binaries; the operational main loop SHALL run ac_tick then bp_tick every invocation.

SW

main/ulp/main.c dispatch #ifdefs

SWR-119

Parking exit limits

BP parking SHALL exit on hall-active, or fail when cumulative encoder ticks reach PARK_ENC_TICK_LIMIT=5200 or wall-clock reaches 30 s (PARK_SAFETY_TIMEOUT_US), reporting bp_hall_detected only on success; determine-parking SHALL publish a bracket only when total ticks lie in [4992, 5376] and SHALL set offset = bracket/2.

SW

spec §5.7.7; main/ulp/main.c parking limits

SWR-120

Park position at hall-zone centre

A calibrated park SHALL stop the BP rotor at the CENTRE of the hall-active (magnet) zone — bracket/2 ticks past hall activation plus any user offset — not at the zone edge.

MECH

parking_position rule; spec §5.7.7

SWR-121

Dose handshake with coast capture

When the CPU arms bp_dosing with bp_dose_target, the ULP SHALL count encoder edges under PID, coast 1 s at target capturing overshoot into bp_dose_actual, then clear bp_dosing and bp_pump_running and wake the CPU.

MECH

spec §5.7.7; main/ulp/main.c dosing

SWR-122

AC bench convergence

With ac_p_step_per_pct=7 the AC pump SHALL converge to a commanded SPR and hold the TRUE seconds-per-revolution — measured as the mean board-side time between consecutive hall revolutions (NOT the encoder-batch sub-sample) — within ±5%, at skip level op_p_skip_max with motor_stop_events = 0.

MECH

spec §5.5 bench table; HIL convergence test (hall-interval SPR)

SWR-123

BP bench convergence

With bp_p_step_per_pct=6 the blood pump SHALL regulate a 16 s/rev target stall-free (motor_stop_events = 0) with the mean hall-interval SPR (board-side time between hall revolutions) within ±5% — the hall interval averages out the within-revolution peristaltic ripple that dominates any single encoder-batch reading.

MECH

spec §5.5 bench table; HIL convergence test (hall-interval SPR)

SWR-124

AC slow-setpoint stability (no limit cycle)

The AC pump SHALL hold SPR 20 without the DAC ever reaching 4095 (DAC_HALT) or re-initializing through the start gate (the historical SPR-20 limit cycle SHALL not recur).

MECH

bug1_ac_spr20 rule; spec §5.5

SWR-125

Live sensor readings on board

On a powered board with emitter rails enabled, the VEML6035 spot (0x3A) and line (0x29) sensors SHALL return non-zero lux and the fluid/pressure ADCs SHALL return in-range readings (0–3.3 V domain).

MECH

spec §4.1/§4.1b; bench sensor probe

3. Main-CPU Firmware (MicroPython)

Persistent storage, hub synchronization, telemetry, the sampling process, sensor thresholds, and firmware safety rules.

ID

Title

Requirement

Class

Source

SWR-201

Append-only store binary format

The event store index SHALL use an 8-byte header (flags byte 0; 24-bit big-endian count at offset 5 when FLAG_LENGTH set, else offset 1), 6-byte entries (uint32 offset + uint16 length), two-phase commit (write inactive slot, sync, flip flag, sync) so an uncommitted append is invisible after reopen, and SHALL reject records > 65535 bytes.

SW

fluispotter/append_only_store.py; test/test_append_only_store.py

SWR-202

Event record JSONL format

EventStore.append SHALL store one newline-terminated compact-JSON record with exactly the keys uid, type, message, arguments, emitted_at, where uid = record index + 0xff0000.

SW

fluispotter/event_store.py; test/test_event_store.py

SWR-203

Event sync upload window

EventStore.synchronize SHALL report the latest local uid, upload only events newer than the server uid as application/jsonl with correct content_length, and perform no upload when empty or already in sync.

SW

fluispotter/event_store.py; test/test_event_store.py

SWR-204

CSV quoting and torn-row recovery

The CSV writer SHALL quote values containing delimiter/terminator/quote (doubling embedded quotes) and on open SHALL position the cursor after the last newline so a power-loss partial row is overwritten.

SW

fluispotter/csv.py; test/test_csv.py

SWR-205

HTTP client protocol behaviors

aiorequests SHALL emit HTTP/1.0 requests with host/connection defaults, Basic auth, JSON body with content-type/length, chunked generator framing, ≤3 redirects (301/302/303 → GET; 307/308 preserve method), and SHALL reject chunked responses with HttpError.

SW

fluispotter/aiorequests.py; test/test_requests.py

SWR-206

Diagnostics CSV schema

diagnostics.csv SHALL carry exactly the 11 columns bp_revs, ac_revs, bp_avg_spr, ac_avg_spr, stalls, fatal_error, fw_version, build_date, board_uid, session_start, session_end — header written only when the file is empty, fatal_error sanitized (',' → ';', newlines stripped).

SW

fluispotter/fluispotter.py _write_diagnostics_row

SWR-207

Time-table binary format

TimeTable entries SHALL persist as 3-byte big-endian '>HB' records (uint16 cumulative minutes, uint8 samples) with exact read/write round-trip and OSError on a partial trailing record.

SW

fluispotter/fluispotter.py TimeTable

SWR-208

Telemetry time-table chunking

Telemetry frames SHALL include tt_idx/tt_n and, when tt_n > 0, a round-robin tt_chunk of min(5, tt_n) (index, minutes, samples) tuples advancing by 5 per frame and resetting when tt_n changes.

SW

fluispotter/fluispotter.py _run_telemetry

SWR-209

Telemetry enablement and cadence

The telemetry task SHALL be gated by the '/telemetry' flag file and emit one JSON object per line every 100 ms, sampling I2C sensors (lux/temp/RH) every 10th frame.

SW

fluispotter/fluispotter.py telemetry constants

SWR-210

Fluid classification thresholds

board_config SHALL define BLOOD_SENSOR_AIR_THRESHOLD_UV=800000, BLOOD_SENSOR_AC_THRESHOLD_UV=500000, BLOOD_SENSOR_BLOOD_THRESHOLD_UV=500000, AC_SENSOR_FLUID_THRESHOLD_UV=500000 (µV; above-air / between-AC / below-blood bands).

SW

fluispotter/board_config.py; spec §4A.6

SWR-211

Dosing pulse targets

The dose steps SHALL command BP_DOSE_TARGET = 720 pulses (PREPARE, 50°) and 3067 pulses (COLLECT, 213°), each equal to round(5184 × degrees / 360).

SW

fluispotter/board_config.py; fluispotter.py _dosing_tick

SWR-212

Sync cadence

The firmware SHALL poll the hub every SYNCHRONIZE_INTERVAL_MS = 1000 ms with a full event-store synchronize every 5th poll.

SW

fluispotter/fluispotter.py _wait_for_command; board_config.py

SWR-213

Sync error containment

A failed synchronize (OSError/HttpError/DeviceError/TimeoutError) SHALL be contained — logged as a Sync-failed event and reported False — never propagating into the device state machine.

SW

fluispotter/fluispotter.py _synchronize_event_store

SWR-214

Event store provisioning

The firmware SHALL create the event store with capacity 8192 backed by events.db.index/events.db and synchronize against http://fluispotter.internal/synchronize.

SW

fluispotter/fluispotter.py constants

SWR-215

No hardware watchdog / no auto-reset

The firmware SHALL NOT instantiate machine.WDT and SHALL NOT auto-reset on ULP-liveness failure: a soft/WDT reset is a false recovery for a latched DAC7574 (operator power-cycle required).

SW

bp_dead_after_crash rule; fluispotter.py

SWR-216

BP direction cleared before RTC I2C handover

ULPPumpRegulator.start() SHALL write the BP direction DAC channel (_set_bp_direction) BEFORE _init_rtc_i2c on every start, so a crash-latched direction register is cleared by a soft reset.

SW

bp_dead_after_crash rule; fluispotter/pump_regulator.py

SWR-217

No blocking stdin in the asyncio loop

fluispotter.py SHALL contain no blocking sys.stdin read inside the asyncio loop (no _run_command_server); any command channel must use a dedicated thread or poll(0) guard.

SW

named_command_channel rule (reverted server)

SWR-218

Blood-detection threshold poll

In CATHETER the firmware SHALL poll the blood receiver every 500 ms and declare blood detected when the reading is below 500000 µV.

SW

fluispotter/fluispotter.py _wait_for_blood_detected

SWR-219

Degree-based movement timeouts

Blood-pump movement timeouts SHALL be computed as int(SPR × 30 × 1000 / 360 × degrees) — 66666 ms for 50°, 284000 ms for 213°, 480000 ms for 360° at SPR 16 (30× safety margin).

SW

spec §5.7.2; fluispotter.py

SWR-220

Dose state machine order

Dosing SHALL execute PARK_1 → ENSURE_CLOSED → PREPARE(720) → OPEN_VALVE → COLLECT(3067) → SMEAR_AND_CLOSE → ADVANCE_TAPE → PARK_2 at a 50 ms tick with the specified per-step timeouts.

SW

spec §5.7.6; fluispotter.py _dosing_tick

SWR-221

Sample confirmation gate

The session loop SHALL transition TRIGGER → DOSING only when all 7 hall-history deque entries post-date the scheduled sample time, and IDLE → TRIGGER when remaining time ≤ 300000 ms.

SW

spec §5.6; fluispotter.py session loop

SWR-222

Fatal-error handling

Each of the 7 fatal movement timeouts SHALL log its specified message to the event store, persist it to fatal-error.txt, and enter the 500 ms LED error-blink state until reed-switch reset.

SW

spec §5.8; fluispotter.py

SWR-223

board_config canonical constants

board_config.py SHALL define the canonical hardware constants: BLOOD_PUMP_SECONDS_PER_REVOLUTION=16, BLOOD_PUMP_START_POWER=3000, AC fast/medium/slow SPR 7/20/20 with start powers 4050/700/700, ULP_PULSES_PER_OUTPUT_REV=5184, PUMP_MOTOR_POWER_HALT=4095, PWM_MAX=65534, PWM_FREQUENCY=2000, VALVE_MOTOR_POWER=100, VEML_SPOT_ADDRESS=0x3a, VEML_LINE_ADDRESS=0x29, SYNCHRONIZE_INTERVAL_MS=1000.

SW

fluispotter/board_config.py; spec §4A.6

SWR-224

Boot config overlay

At boot the firmware SHALL load a tunable-calibration overlay from /board_overrides.json over the compiled board_config defaults, applied once (no runtime mutation); a missing or malformed overlay SHALL degrade to the compiled defaults without raising.

SW

fluispotter/config.py; fluispotter.py BaseController.create()

SWR-225

Overlay allowlist + hardware guard

The overlay SHALL accept ONLY the 25 tunable calibration keys (config.TUNABLE_KEYS); any hardware key (PIN_*, DAC_REGISTER_*, VEML_*, PUMP_MOTOR_*, PWM_*, VALVE_MOTOR_*, ULP_PULSES_PER_OUTPUT_REV), unknown key, non-int, or out-of-range value SHALL be rejected (kept at default) and recorded.

SW

fluispotter/config.py Config._validate

SWR-226

Config hash

The effective config SHALL be fingerprinted with FNV-1a 32-bit over a sorted-key serialization (deterministic, order-independent), identical to output/_serial_io.fnv, emitted as 8 hex digits.

SW

fluispotter/config.py fnv1a/_compute_hash

SWR-227

Boot config event + change detection

On boot the firmware SHALL append one event_store record of type 'config' carrying the hash, source (defaults/overlay/locked), override count and rejected list, and SHALL set a changed flag by comparing the boot hash to the last persisted hash (/board_config.hash), then persist the new hash.

SW

fluispotter.py create(); fluispotter/config.py

SWR-228

Config lock (freeze)

When /board_config.lock is present the firmware SHALL ignore the overlay and use the compiled defaults only (source=locked) — the hook by which calibration can be frozen back into firmware.

SW

fluispotter/config.py _is_locked

SWR-229

Tunable consumers route through the accessor

Every tunable-value consumer in fluispotter.py SHALL read through the boot-loaded Config accessor (self.cfg.NAME) rather than the bare compiled constant, so the overlay takes effect; hardware constants keep the compiled value.

SW

fluispotter/fluispotter.py

4. Desktop Applications (Monitor & Control)

The macOS apps: the Monitor read-only guarantee, Control mode gating and guards, and the injected MicroPython script contracts.

ID

Title

Requirement

Class

Source

SWR-301

Monitor app is read-only

The Monitor app's SerialBridge SHALL expose no board-write path (no inject/run_command and no serial write call): connecting the Monitor sends zero bytes to the board.

SW

monitor/serial_bridge.py; telemetry-serial rule

SWR-302

App version in window title

Each app SHALL define APP_VERSION matching ^\d+\.\d+\.\d+$ in its app_version.py, list that version in its changelog, and include 'v{APP_VERSION}' in the window title ahead of the git build tag.

SW

app_version_discipline rule; monitor/control main.py

SWR-303

No hardcoded ULP addresses

No source file outside the generated variables modules SHALL contain a hardcoded ULP RTC address literal (0x5000xxxx); consumers SHALL import addresses by name.

SW

ulp_deploy_checklist rule; board_script/control

SWR-304

Injected scripts compile

MONITOR_SCRIPT and the output of every board_script factory SHALL compile (compile(src, ..., 'exec')) for default arguments and both pump='ac' and pump='bp'.

SW

monitor/board_script.py

SWR-305

MONITOR_SCRIPT frame contract

The MONITOR_SCRIPT telemetry frame SHALL contain the documented keys (t, inv, main_state, ac_*/bp_* regulator fields, actuators, sensors, schedule) with ac_pwr/bp_pwr computed as 4095 − dac.

SW

monitor/board_script.py MONITOR_SCRIPT

SWR-306

Control connect-time address guard

On connect the Control app SHALL probe ULP_VERSION and INITIALIZED (twice) at its loaded addresses and refuse the connection unless the packed version is sane (major == 1) and INITIALIZED is advancing.

SW

control/main.py _verify_board_addresses; stale-address rule

SWR-307

Control mode gating

The Control app SHALL run at most one injected script at a time, route frames by namespace key and mode, and on any done/err frame revert to monitor mode re-injecting MONITOR_SCRIPT.

SW

control/main.py _handle_frame

SWR-308

Frame history and Sensors repopulate

The Control app SHALL append every received frame to a 4000-deep history and rebuild the Sensors sparklines from it when the tab is restored, accepting readings at frame top level and under the sensor key (full continuity across tab switches).

SW

control/main.py; control/panels/sensors_panel.py repopulate

SWR-309

Fluid-sensor state-bounds classification

The Control Sensors tab SHALL classify the fluid sensors with independent per-state [lower, upper] bounds in the signal ordering Blood < Air < Water/AC (readings between bounds classify as TRANSITION), and SHALL expose the four directed transition values — Air→AC, AC→Blood, Blood→Air, Air→Blood — as host-side tunables evaluated live against the retrospective buffer.

SW

control/panels/sensors_panel.py _CLASSIFY (states model)

SWR-310

Board Config tab parses source

Both apps' Board Config tab SHALL render board_config.py by parsing its source text (never executing it), preserving section grouping and order, currently 68 constants in 11 sections.

SW

monitor+control panels/board_config_panel.py

SWR-311

VEML read protocol in host scripts

Every VEML6035 reader in the injected scripts SHALL write CONFIG 0x00 = 0x0300 (power-on) before reading ALS data at register 0x04 and scale lux = count × 0.0512.

SW

spec §4.1; monitor/board_script.py

SWR-312

ADC full-range attenuation

Every ADC reader for the fluid (GPIO 6/7) and pressure (GPIO 3) channels SHALL pass atten=ADC.ATTN_11DB for the full 0–3.3 V range.

SW

spec §4.1b; firmware + board_script ADC constructors

SWR-313

park_pump_script guards

park_pump_script SHALL refuse pump != 'bp', abort when bp_dosing/bp_pump_running/bp_parking is set, and write the locked profile (target 16000 ms, Kp=32/Ki=0/Kd=0, DAC register 895) before arming bp_parking.

SW

monitor/board_script.py park_pump_script

SWR-314

Sensors stream during PID runs

All three PID-tuner run scripts SHALL emit {sensor} frames from a fully exception-guarded sensor block so the Sensors tab stays live during a run without perturbing the regulation.

SW

monitor/board_script.py _LIVE_SENSORS_SRC

5. Synchronize Service (RPi Hub)

The hub HTTP API contract the device and tools depend on.

ID

Title

Requirement

Class

Source

SWR-401

Hub REST endpoint set

The synchronize service SHALL listen on port 9090 and expose the documented device/event/file/settings endpoints (GET/POST devices, paginated events, batch event upload, streamed files, settings).

SW

spec §6.3; lib/synchronize-service

SWR-402

Registration and command acknowledgement

POST /devices/{uid} SHALL create-or-update and return settings (time table + pending commands); an uploaded command event citing a command_uid SHALL clear that pending command.

SW

spec §6.5; lib/synchronize-service

SWR-403

Device-field validation

The service SHALL validate firmware_version against ^\d+\.\d+\.\d+$, mac_address as exactly 6 bytes, ip_address as 4 or 16 bytes.

SW

spec §6.8; lib/synchronize-service

SWR-404

Content-addressed file storage

Uploaded device files SHALL be stored content-addressed by SHA-256 with byte-identical blob retrieval.

SW

spec §6.3-6.5; lib/synchronize-service

6. Build, Deploy & Process Gates

Build-time verification gates, deploy verification, and the version disciplines enforced in CI.

ID

Title

Requirement

Class

Source

SWR-501

ULP build gates

Every ULP build SHALL pass: stack headroom ≥ 768 bytes (__stack_top − end of .bss), no new libgcc intrinsics beyond __ashldi3, runtime-loaded test binaries ≤ 4592 bytes, operational binary < 8192 bytes.

SW

ulp_deploy_checklist; output/build_verify_ulp.py

SWR-502

Address maps importable and complete

The generated variables.py / test_variables.py / test2_variables.py SHALL import standalone on desktop Python and expose the required shared-variable names including the Simple-P knobs (OP_P_*, *_P_STEP_PER_PCT) and ULP_VERSION.

SW

fluispotter/variables*.py; deploy checklist

SWR-503

Binary discriminator symbols

test_determine_parking SHALL exist only in the operational binary and test_verify_encoder only in the test binary, so a loaded binary is uniquely identifiable.

SW

multi_mode_ulp; ELF symbol check

SWR-504

Full-binary deploy verification

A ULP deploy SHALL byte-compare the FULL binary against RTC slow memory and abort on the first mismatch (a reset-vector check is insufficient).

MECH

ulp_deploy_checklist step 11

SWR-505

Post-deploy telemetry sanity

After a deploy + reset, telemetry SHALL report the freshly built ulp_ver, an advancing INITIALIZED counter, and in-range regulator fields (dac 0-4095, states 0-6).

MECH

ulp_deploy_checklist

SWR-506

Post-deploy motion proof

After a deploy, arming BP_PUMP_RUNNING SHALL produce encoder pulses and hall-revolution increments within 2 s (state-word advance alone is not proof of motion).

MECH

ulp_deploy_checklist

SWR-507

Binary-swap self-test

A deploy SHALL pass the three-phase swap test: operational parking calibration → test-binary encoder check → operational parking calibration again, aborting on any phase failure.

MECH

ulp_deploy_checklist step 14

SWR-508

Version-bump discipline

Any change to monitor/control app code SHALL bump that app's APP_VERSION (board_script.py counts as Control), and any change to the ULP main.c SHALL bump a ULP_VERSION_* line — enforced as a CI gate over the PR range.

SW

app_version_discipline; .github/scripts/check_version_bump.py

SWR-509

Spec-document version consistency

Each living spec document SHALL carry the same version in its cover line and its document-control/overview table, and that version SHALL appear in its change log.

SW

spec-repo update rules; docs/*.html

SWR-510

Requirements doc mirrors the registry

docs/fluispotter_sw_requirements.html SHALL contain every requirement ID in this registry and no unknown SWR- IDs (generated doc, drift check).

SW

verification/requirements.py; gen_requirements_doc.py

7. Verification & Traceability

The verification framework (verification/) executes the requirement-tagged test suites and emits a per-requirement PASS / FAIL / SKIP / NOT-TESTED report including the version of every SW component (git build, ULP ×3, Monitor/Control app versions, spec-document versions, synchronize-service SHA):

python3 -m verification.run --suite sw — the headless SW suite; runs on every build in Stage 0 CI.

      python3 -m verification.run --suite mech --board — the mechanics-dependent suite; bench/HIL stage only (drives motors, requires the connected board and the explicit --board flag — without it every MECH test is skipped).

Requirements with no referencing test in the executed suite are reported NOT-TESTED, so coverage gaps are always visible in the report rather than silently green.

— End of Fluispotter SW Requirements, Version 1.5 (generated from verification/requirements.py) —